## Information
Laser Diode Information
Datasheet Information
Brand: | Roithner Lasers
Part Number: | S9850MG
Package Size: | 5.6 mm (TO56), Flat window
Photodiode: | Integrated
Wavelength: | 980 | nm
Optical Power: | 50 | mW
Threshold Current: | 12 | mA
Operating Current: | 75 | mA
Slope Efficiency: | 0.8 | W/A
Beam Divergence (FWHM)
parallel: | 10 | deg
perpendicular: | 35 | deg

## Angle_Parallel
Optical Power Output vs Angle |  |  | Parallel Direction
Measurment Information: |  |  | Measured Data: |  |  |  |  | Calculated Data
|  |  | PSA = Polarization State Analyser |  |  |  |  | * DOLP - degree of linear polarization
PowerMeter: | Thorlabs PM320E
PowerMeter Detector: | Thorlabs S142C |  | Angle | Intensity | Intensity | Intensity | Intensity | Stocks Vector Components |  |  | Norm. Stockes Vector Comp. |  |  | DOLP* | Azimuth
Laser Diode Controller: | Thorlabs LDC205C |  | (⁰) | (a.u.) | (a.u.) | (a.u.) | (a.u.) | s0 | s1 | s2 | S0 | S1 | S2 |  | (⁰)
Wave Generator: | Teledyne LeCroy WaveStation 2023 |  |  | PSA = 0⁰ | PSA = 45⁰ | PSA = 90⁰ | PSA = 135⁰
Temperature Controller: | Thorlabs TED200C |  | -60 | 2.0406e-05 | 1.8808e-05 | 2.5144e-05 | 2.8499e-05 | 4.555e-05 | -4.738e-06 | -9.691e-06 | 1 | -0.10402 | -0.21276 | 0.23682 | 0.55803
Laser Diode Holder: | Thorlabs TCLDM9 |  | -59.5 | 2.6941e-05 | 2.2523e-05 | 2.512e-05 | 2.6457e-05 | 5.2061e-05 | 1.821e-06 | -3.934e-06 | 1 | 0.034978 | -0.075565 | 0.083268 | -0.56864
Detector | Hamamatsu H9656-20 |  | -59 | 2.5187e-05 | 2.4229e-05 | 3.0404e-05 | 3.2092e-05 | 5.5591e-05 | -5.217e-06 | -7.863e-06 | 1 | -0.093846 | -0.14144 | 0.16975 | 0.4925
Lock-in Amplifier | Stanford Research System SR830 |  | -58.5 | 2.1853e-05 | 2.2046e-05 | 2.4242e-05 | 2.3734e-05 | 4.6095e-05 | -2.389e-06 | -1.688e-06 | 1 | -0.051828 | -0.03662 | 0.06346 | 0.30756
Measurment Notes |  |  | -58 | 3.1837e-05 | 2.7999e-05 | 3.6054e-05 | 3.8268e-05 | 6.7891e-05 | -4.217e-06 | -1.0269e-05 | 1 | -0.062114 | -0.15126 | 0.16351 | 0.59057
|  |  | -57.5 | 3.0544e-05 | 2.6692e-05 | 3.3405e-05 | 3.7203e-05 | 6.3949e-05 | -2.861e-06 | -1.0511e-05 | 1 | -0.044739 | -0.16437 | 0.17035 | 0.65252
Fixed LD house Temp.: | 20 | ⁰C | -57 | 3.3099e-05 | 2.4654e-05 | 2.1255e-05 | 3.0482e-05 | 5.4354e-05 | 1.1844e-05 | -5.828e-06 | 1 | 0.2179 | -0.10722 | 0.24286 | -0.22864
Optical Power (avg.): | 4.9975 | mW | -56.5 | 3.8378e-05 | 2.9449e-05 | 2.747e-05 | 3.7505e-05 | 6.5848e-05 | 1.0908e-05 | -8.056e-06 | 1 | 0.16565 | -0.12234 | 0.20593 | -0.31806
LD Photodiode Current: | 0.016 | mA | -56 | 3.4265e-05 | 2.6684e-05 | 2.458e-05 | 3.1756e-05 | 5.8845e-05 | 9.685e-06 | -5.072e-06 | 1 | 0.16458 | -0.086193 | 0.18579 | -0.24121
Modulation Freq.: | 275 | Hz | -55.5 | 2.7053e-05 | 2.0776e-05 | 1.7556e-05 | 2.1666e-05 | 4.4609e-05 | 9.497e-06 | -8.9e-07 | 1 | 0.21289 | -0.019951 | 0.21383 | -0.04672
|  |  | -55 | 4.1932e-05 | 2.9609e-05 | 2.3421e-05 | 3.4027e-05 | 6.5353e-05 | 1.8511e-05 | -4.418e-06 | 1 | 0.28325 | -0.067602 | 0.2912 | -0.11714
FWHM | 7.9649 | ⁰ | -54.5 | 3.1531e-05 | 2.4863e-05 | 2.3948e-05 | 2.8962e-05 | 5.5479e-05 | 7.583e-06 | -4.099e-06 | 1 | 0.13668 | -0.073884 | 0.15537 | -0.24778
DOLP (max value) | 98.555 | % | -54 | 3.2656e-05 | 2.1525e-05 | 2.3637e-05 | 3.771e-05 | 5.6293e-05 | 9.019e-06 | -1.6185e-05 | 1 | 0.16022 | -0.28751 | 0.32914 | -0.5312
|  |  | -53.5 | 3.6374e-05 | 2.4095e-05 | 2.7253e-05 | 4.0487e-05 | 6.3627e-05 | 9.121e-06 | -1.6392e-05 | 1 | 0.14335 | -0.25763 | 0.29482 | -0.53151
Auxiliar |  |  | -53 | 3.6054e-05 | 2.2907e-05 | 3.1671e-05 | 4.5099e-05 | 6.7725e-05 | 4.383e-06 | -2.2192e-05 | 1 | 0.064718 | -0.32768 | 0.33401 | -0.6879
Max. Intentity PSA = 0º | 0.62055 |  | -52.5 | 2.4289e-05 | 2.1488e-05 | 1.9228e-05 | 2.1953e-05 | 4.3517e-05 | 5.061e-06 | -4.65e-07 | 1 | 0.1163 | -0.010685 | 0.11679 | -0.045811
Row Position = Max. | 127 |  | -52 | 2.3354e-05 | 1.7133e-05 | 1.7339e-05 | 2.3285e-05 | 4.0693e-05 | 6.015e-06 | -6.152e-06 | 1 | 0.14781 | -0.15118 | 0.21143 | -0.39833
Half-Maximum | 0.31027 |  | -51.5 | 2.2233e-05 | 1.8053e-05 | 2.069e-05 | 2.4745e-05 | 4.2923e-05 | 1.543e-06 | -6.692e-06 | 1 | 0.035948 | -0.15591 | 0.16 | -0.67209
Negative 'Half-angle' | -4.9161 | ⁰ | -51 | 2.4278e-05 | 2.2229e-05 | 2.1646e-05 | 2.4067e-05 | 4.5924e-05 | 2.632e-06 | -1.838e-06 | 1 | 0.057312 | -0.040023 | 0.069903 | -0.3048
Positive 'Half-angle' | 3.0489 | ⁰ | -50.5 | 2.1242e-05 | 1.714e-05 | 1.5313e-05 | 1.9027e-05 | 3.6555e-05 | 5.929e-06 | -1.887e-06 | 1 | 0.16219 | -0.051621 | 0.17021 | -0.15406
|  |  | -50 | 2.0899e-05 | 1.7591e-05 | 1.6781e-05 | 1.9893e-05 | 3.768e-05 | 4.118e-06 | -2.302e-06 | 1 | 0.10929 | -0.061093 | 0.12521 | -0.25487
|  |  | -49.5 | 2.3939e-05 | 1.961e-05 | 1.9903e-05 | 2.3978e-05 | 4.3842e-05 | 4.036e-06 | -4.368e-06 | 1 | 0.092058 | -0.09963 | 0.13565 | -0.41244
|  |  | -49 | 2.0527e-05 | 1.8604e-05 | 1.5246e-05 | 1.7159e-05 | 3.5773e-05 | 5.281e-06 | 1.445e-06 | 1 | 0.14763 | 0.040394 | 0.15305 | 0.13354
|  |  | -48.5 | 2.3212e-05 | 1.8552e-05 | 1.5302e-05 | 1.9919e-05 | 3.8514e-05 | 7.91e-06 | -1.367e-06 | 1 | 0.20538 | -0.035494 | 0.20842 | -0.085564
|  |  | -48 | 2.5746e-05 | 2.1223e-05 | 2.1294e-05 | 2.6584e-05 | 4.704e-05 | 4.452e-06 | -5.361e-06 | 1 | 0.094643 | -0.11397 | 0.14814 | -0.43888
|  |  | -47.5 | 2.395e-05 | 2.1599e-05 | 1.9984e-05 | 2.268e-05 | 4.3934e-05 | 3.966e-06 | -1.081e-06 | 1 | 0.090272 | -0.024605 | 0.093565 | -0.13305
|  |  | -47 | 2.7128e-05 | 2.015e-05 | 1.9575e-05 | 2.6396e-05 | 4.6703e-05 | 7.553e-06 | -6.246e-06 | 1 | 0.16172 | -0.13374 | 0.20986 | -0.34548
|  |  | -46.5 | 2.9679e-05 | 2.9423e-05 | 2.899e-05 | 3.0097e-05 | 5.8669e-05 | 6.89e-07 | -6.74e-07 | 1 | 0.011744 | -0.011488 | 0.016429 | -0.3872
|  |  | -46 | 3.0164e-05 | 2.6211e-05 | 2.248e-05 | 2.6673e-05 | 5.2644e-05 | 7.684e-06 | -4.62e-07 | 1 | 0.14596 | -0.0087759 | 0.14623 | -0.030026
|  |  | -45.5 | 3.0842e-05 | 2.6245e-05 | 2.067e-05 | 2.5798e-05 | 5.1512e-05 | 1.0172e-05 | 4.47e-07 | 1 | 0.19747 | 0.0086776 | 0.19766 | 0.021958
|  |  | -45 | 3.6199e-05 | 2.9132e-05 | 2.685e-05 | 3.4217e-05 | 6.3049e-05 | 9.349e-06 | -5.085e-06 | 1 | 0.14828 | -0.080652 | 0.1688 | -0.24908
|  |  | -44.5 | 4.6298e-05 | 3.3062e-05 | 2.3581e-05 | 3.6085e-05 | 6.9879e-05 | 2.2717e-05 | -3.023e-06 | 1 | 0.32509 | -0.04326 | 0.32796 | -0.066147
|  |  | -44 | 7.0923e-05 | 4.6384e-05 | 2.4768e-05 | 4.999e-05 | 9.5691e-05 | 4.6155e-05 | -3.606e-06 | 1 | 0.48233 | -0.037684 | 0.4838 | -0.038985
|  |  | -43.5 | 0.00011295 | 6.9324e-05 | 3.4617e-05 | 7.7278e-05 | 0.00014757 | 7.8334e-05 | -7.954e-06 | 1 | 0.53083 | -0.053901 | 0.53356 | -0.050596
|  |  | -43 | 0.00016841 | 0.00010532 | 4.9254e-05 | 0.00011445 | 0.00021767 | 0.00011916 | -9.127e-06 | 1 | 0.54744 | -0.041931 | 0.54904 | -0.038223
|  |  | -42.5 | 0.00019902 | 0.00011692 | 4.6149e-05 | 0.00012643 | 0.00024517 | 0.00015287 | -9.515e-06 | 1 | 0.62353 | -0.03881 | 0.62474 | -0.031081
|  |  | -42 | 0.00019175 | 0.00011042 | 4.494e-05 | 0.00012411 | 0.00023669 | 0.00014681 | -1.3695e-05 | 1 | 0.62026 | -0.057861 | 0.62295 | -0.046508
|  |  | -41.5 | 0.00019532 | 0.00011677 | 4.9439e-05 | 0.00012693 | 0.00024476 | 0.00014589 | -1.0163e-05 | 1 | 0.59603 | -0.041522 | 0.59747 | -0.034776
|  |  | -41 | 0.00019437 | 0.00011176 | 5.0627e-05 | 0.00013141 | 0.000245 | 0.00014375 | -1.9648e-05 | 1 | 0.58672 | -0.080196 | 0.59217 | -0.067922
|  |  | -40.5 | 0.0001857 | 0.00010681 | 4.8213e-05 | 0.00012472 | 0.00023391 | 0.00013749 | -1.7912e-05 | 1 | 0.58777 | -0.076576 | 0.59274 | -0.064776
|  |  | -40 | 0.00019562 | 0.00011194 | 4.9144e-05 | 0.00013121 | 0.00024477 | 0.00014648 | -1.9268e-05 | 1 | 0.59844 | -0.07872 | 0.6036 | -0.065395
|  |  | -39.5 | 0.0002121 | 0.00012353 | 5.1171e-05 | 0.00013834 | 0.00026327 | 0.00016093 | -1.4811e-05 | 1 | 0.61127 | -0.056257 | 0.61386 | -0.045887
|  |  | -39 | 0.00025091 | 0.00014508 | 5.3495e-05 | 0.00015789 | 0.0003044 | 0.00019741 | -1.2807e-05 | 1 | 0.64853 | -0.042073 | 0.64989 | -0.032392
|  |  | -38.5 | 0.00028038 | 0.0001601 | 5.6639e-05 | 0.0001732 | 0.00033702 | 0.00022374 | -1.3098e-05 | 1 | 0.66388 | -0.038864 | 0.66502 | -0.029237
|  |  | -38 | 0.00027338 | 0.00016266 | 5.8793e-05 | 0.00016652 | 0.00033217 | 0.00021459 | -3.859e-06 | 1 | 0.64601 | -0.011617 | 0.64611 | -0.0089908
|  |  | -37.5 | 0.0002805 | 0.00015813 | 6.3524e-05 | 0.0001813 | 0.00034402 | 0.00021698 | -2.3171e-05 | 1 | 0.6307 | -0.067353 | 0.63429 | -0.053193
|  |  | -37 | 0.00028849 | 0.00016323 | 5.7541e-05 | 0.00018227 | 0.00034603 | 0.00023095 | -1.9044e-05 | 1 | 0.66742 | -0.055036 | 0.66969 | -0.041137
|  |  | -36.5 | 0.00030747 | 0.00017947 | 7.1355e-05 | 0.00020653 | 0.00037883 | 0.00023612 | -2.706e-05 | 1 | 0.62328 | -0.071431 | 0.62736 | -0.057053
|  |  | -36 | 0.00030273 | 0.00017649 | 6.0059e-05 | 0.00018591 | 0.00036279 | 0.00024267 | -9.418e-06 | 1 | 0.66891 | -0.02596 | 0.66941 | -0.019395
|  |  | -35.5 | 0.00034318 | 0.0001992 | 6.1512e-05 | 0.00019836 | 0.00040469 | 0.00028166 | 8.35e-07 | 1 | 0.696 | 0.0020633 | 0.69601 | 0.0014823
|  |  | -35 | 0.00034628 | 0.00020734 | 6.6393e-05 | 0.00020552 | 0.00041267 | 0.00027988 | 1.818e-06 | 1 | 0.67823 | 0.0044055 | 0.67824 | 0.0032478
|  |  | -34.5 | 0.00033435 | 0.00019401 | 6.1036e-05 | 0.00020033 | 0.00039539 | 0.00027332 | -6.318e-06 | 1 | 0.69126 | -0.015979 | 0.69145 | -0.011556
|  |  | -34 | 0.0003686 | 0.00022441 | 7.3061e-05 | 0.00021553 | 0.00044166 | 0.00029554 | 8.881e-06 | 1 | 0.66915 | 0.020108 | 0.66945 | 0.015021
|  |  | -33.5 | 0.0003357 | 0.0002088 | 7.2346e-05 | 0.00019956 | 0.00040804 | 0.00026335 | 9.239e-06 | 1 | 0.6454 | 0.022642 | 0.6458 | 0.017534
|  |  | -33 | 0.0003649 | 0.00022295 | 7.1832e-05 | 0.00021237 | 0.00043673 | 0.00029307 | 1.058e-05 | 1 | 0.67105 | 0.024225 | 0.67149 | 0.018043
|  |  | -32.5 | 0.00035221 | 0.00022405 | 9.209e-05 | 0.00022149 | 0.0004443 | 0.00026012 | 2.563e-06 | 1 | 0.58546 | 0.0057687 | 0.58548 | 0.0049265
|  |  | -32 | 0.00037751 | 0.00021565 | 7.5169e-05 | 0.00023788 | 0.00045268 | 0.00030234 | -2.2232e-05 | 1 | 0.66789 | -0.049112 | 0.66969 | -0.036701
|  |  | -31.5 | 0.00032881 | 0.00021145 | 7.5169e-05 | 0.00019062 | 0.00040398 | 0.00025364 | 2.0831e-05 | 1 | 0.62786 | 0.051564 | 0.62997 | 0.040972
|  |  | -31 | 0.00031737 | 0.00019851 | 7.7561e-05 | 0.0001955 | 0.00039493 | 0.00023981 | 3.01e-06 | 1 | 0.60721 | 0.0076216 | 0.60726 | 0.0062756
|  |  | -30.5 | 0.00034261 | 0.00019732 | 7.2532e-05 | 0.00021434 | 0.00041514 | 0.00027008 | -1.7017e-05 | 1 | 0.65057 | -0.040991 | 0.65186 | -0.031462
|  |  | -30 | 0.00033197 | 0.00020501 | 7.5661e-05 | 0.00020045 | 0.00040763 | 0.00025631 | 4.56e-06 | 1 | 0.62878 | 0.011187 | 0.62888 | 0.0088946
|  |  | -29.5 | 0.00032917 | 0.0001986 | 8.189e-05 | 0.00021178 | 0.00041106 | 0.00024728 | -1.3172e-05 | 1 | 0.60157 | -0.032044 | 0.60242 | -0.026609
|  |  | -29 | 0.00031969 | 0.00018963 | 7.3441e-05 | 0.00020051 | 0.00039313 | 0.00024625 | -1.0878e-05 | 1 | 0.62638 | -0.02767 | 0.62699 | -0.022073
|  |  | -28.5 | 0.00033692 | 0.00020349 | 8.3201e-05 | 0.00021684 | 0.00042012 | 0.00025372 | -1.3352e-05 | 1 | 0.60392 | -0.031782 | 0.60475 | -0.026289
|  |  | -28 | 0.00033784 | 0.00019181 | 8.6725e-05 | 0.00023043 | 0.00042457 | 0.00025112 | -3.8623e-05 | 1 | 0.59147 | -0.090971 | 0.59842 | -0.076305
|  |  | -27.5 | 0.00034678 | 0.00019947 | 8.7507e-05 | 0.00022703 | 0.00043429 | 0.00025928 | -2.7567e-05 | 1 | 0.59701 | -0.063476 | 0.60038 | -0.052963
|  |  | -27 | 0.00035563 | 0.00019214 | 8.8022e-05 | 0.00024688 | 0.00044365 | 0.00026761 | -5.4747e-05 | 1 | 0.6032 | -0.1234 | 0.61569 | -0.1009
|  |  | -26.5 | 0.00036696 | 0.00020164 | 8.7984e-05 | 0.00025481 | 0.00045494 | 0.00027897 | -5.3169e-05 | 1 | 0.61321 | -0.11687 | 0.62425 | -0.094165
|  |  | -26 | 0.00039637 | 0.00020549 | 8.8327e-05 | 0.00029224 | 0.0004847 | 0.00030805 | -8.6755e-05 | 1 | 0.63554 | -0.17899 | 0.66026 | -0.13726
|  |  | -25.5 | 0.0003936 | 0.00021106 | 0.00010958 | 0.00030047 | 0.00050318 | 0.00028402 | -8.9408e-05 | 1 | 0.56446 | -0.17769 | 0.59177 | -0.15248
|  |  | -25 | 0.00047449 | 0.0002462 | 0.00010386 | 0.00031603 | 0.00057835 | 0.00037062 | -6.9827e-05 | 1 | 0.64083 | -0.12074 | 0.65211 | -0.093111
|  |  | -24.5 | 0.00063718 | 0.00028321 | 0.00010297 | 0.00046075 | 0.00074015 | 0.0005342 | -0.00017753 | 1 | 0.72175 | -0.23986 | 0.76056 | -0.16043
|  |  | -24 | 0.00058163 | 0.00027648 | 0.00010134 | 0.00041676 | 0.00068297 | 0.00048028 | -0.00014028 | 1 | 0.70323 | -0.2054 | 0.73261 | -0.14209
|  |  | -23.5 | 0.00069428 | 0.00031817 | 0.00011158 | 0.00046027 | 0.00080586 | 0.0005827 | -0.0001421 | 1 | 0.72308 | -0.17633 | 0.74427 | -0.1196
|  |  | -23 | 0.00062812 | 0.00031981 | 0.00011448 | 0.00042761 | 0.0007426 | 0.00051364 | -0.0001078 | 1 | 0.69168 | -0.14516 | 0.70675 | -0.10343
|  |  | -22.5 | 0.00072229 | 0.00029797 | 0.00013825 | 0.00056643 | 0.00086054 | 0.00058405 | -0.00026846 | 1 | 0.6787 | -0.31197 | 0.74697 | -0.21543
|  |  | -22 | 0.00060058 | 0.00028756 | 0.00011343 | 0.00043899 | 0.00071401 | 0.00048715 | -0.00015143 | 1 | 0.68228 | -0.21208 | 0.71448 | -0.15069
|  |  | -21.5 | 0.00081587 | 0.00035799 | 0.00012842 | 0.00061506 | 0.00094429 | 0.00068745 | -0.00025708 | 1 | 0.72801 | -0.27224 | 0.77725 | -0.17893
|  |  | -21 | 0.00063932 | 0.00029454 | 0.00013471 | 0.00045824 | 0.00077403 | 0.00050462 | -0.0001637 | 1 | 0.65193 | -0.2115 | 0.68538 | -0.15685
|  |  | -20.5 | 0.00055123 | 0.00030193 | 0.00021535 | 0.00045055 | 0.00076658 | 0.00033587 | -0.00014862 | 1 | 0.43815 | -0.19388 | 0.47913 | -0.2083
|  |  | -20 | 0.00055826 | 0.00027132 | 0.00012964 | 0.00042921 | 0.0006879 | 0.00042862 | -0.00015789 | 1 | 0.62308 | -0.22953 | 0.66401 | -0.17648
|  |  | -19.5 | 0.00078917 | 0.00034252 | 0.0001331 | 0.00056273 | 0.00092227 | 0.00065607 | -0.00022021 | 1 | 0.71137 | -0.23877 | 0.75037 | -0.16192
|  |  | -19 | 0.00063002 | 0.00028897 | 0.00015223 | 0.00047547 | 0.00078226 | 0.00047779 | -0.0001865 | 1 | 0.61079 | -0.23842 | 0.65567 | -0.18608
|  |  | -18.5 | 0.00061965 | 0.00034216 | 0.00015038 | 0.00043613 | 0.00077004 | 0.00046927 | -9.3967e-05 | 1 | 0.60941 | -0.12203 | 0.62151 | -0.098813
|  |  | -18 | 0.00066555 | 0.00030872 | 0.00015164 | 0.00049168 | 0.00081718 | 0.00051391 | -0.00018296 | 1 | 0.62888 | -0.22389 | 0.66755 | -0.17101
|  |  | -17.5 | 0.00061894 | 0.00033304 | 0.00014651 | 0.00042612 | 0.00076545 | 0.00047243 | -9.3072e-05 | 1 | 0.61719 | -0.12159 | 0.62906 | -0.097258
|  |  | -17 | 0.00054026 | 0.00038451 | 0.0003081 | 0.00046533 | 0.00084836 | 0.00023216 | -8.0824e-05 | 1 | 0.27366 | -0.095271 | 0.28977 | -0.16751
|  |  | -16.5 | 0.00065053 | 0.00062097 | 0.00034809 | 0.00039643 | 0.00099862 | 0.00030244 | 0.00022453 | 1 | 0.30285 | 0.22484 | 0.37719 | 0.31931
|  |  | -16 | 0.0011294 | 0.00068665 | 0.00025249 | 0.00069362 | 0.0013819 | 0.00087691 | -6.974e-06 | 1 | 0.63458 | -0.0050467 | 0.6346 | -0.0039764
|  |  | -15.5 | 0.0012792 | 0.00073648 | 0.00017983 | 0.00070268 | 0.0014591 | 0.0010994 | 3.3796e-05 | 1 | 0.7535 | 0.023163 | 0.75386 | 0.015365
|  |  | -15 | 0.0015109 | 0.00084472 | 0.00031614 | 0.001029 | 0.001827 | 0.0011947 | -0.0001843 | 1 | 0.65392 | -0.10087 | 0.66166 | -0.076527
|  |  | -14.5 | 0.0020881 | 0.0012275 | 0.00029975 | 0.0010958 | 0.0023878 | 0.0017883 | 0.00013173 | 1 | 0.74893 | 0.055167 | 0.75096 | 0.036764
|  |  | -14 | 0.003707 | 0.0022874 | 0.00046092 | 0.0020094 | 0.0041679 | 0.003246 | 0.00027799 | 1 | 0.77882 | 0.066698 | 0.78167 | 0.042716
|  |  | -13.5 | 0.0061989 | 0.0039468 | 0.00059635 | 0.0027206 | 0.0067953 | 0.0056026 | 0.0012262 | 1 | 0.82448 | 0.18045 | 0.844 | 0.10773
|  |  | -13 | 0.0091057 | 0.0053606 | 0.00048286 | 0.0042906 | 0.0095886 | 0.0086229 | 0.00107 | 1 | 0.89928 | 0.11159 | 0.90618 | 0.06173
|  |  | -12.5 | 0.01467 | 0.0076375 | 0.00057292 | 0.0077238 | 0.015242 | 0.014097 | -8.631e-05 | 1 | 0.92483 | -0.0056625 | 0.92484 | -0.0030613
|  |  | -12 | 0.01346 | 0.006804 | 0.00068707 | 0.0072871 | 0.014147 | 0.012773 | -0.00048304 | 1 | 0.90287 | -0.034144 | 0.90351 | -0.018899
|  |  | -11.5 | 0.018213 | 0.0097628 | 0.00095523 | 0.0094567 | 0.019169 | 0.017258 | 0.00030613 | 1 | 0.90033 | 0.01597 | 0.90048 | 0.0088682
|  |  | -11 | 0.028992 | 0.016331 | 0.0013194 | 0.014033 | 0.030311 | 0.027672 | 0.0022974 | 1 | 0.91294 | 0.075794 | 0.91608 | 0.041416
|  |  | -10.5 | 0.039715 | 0.021931 | 0.0012074 | 0.019418 | 0.040922 | 0.038508 | 0.0025129 | 1 | 0.94099 | 0.061406 | 0.94299 | 0.032582
|  |  | -10 | 0.055664 | 0.028633 | 0.0013726 | 0.028835 | 0.057037 | 0.054292 | -0.0002022 | 1 | 0.95187 | -0.0035451 | 0.95188 | -0.0018622
|  |  | -9.5 | 0.074311 | 0.036839 | 0.0016377 | 0.039681 | 0.075949 | 0.072673 | -0.0028419 | 1 | 0.95687 | -0.037419 | 0.9576 | -0.019543
|  |  | -9 | 0.090935 | 0.044619 | 0.0020147 | 0.049618 | 0.09295 | 0.088921 | -0.0049992 | 1 | 0.95665 | -0.053784 | 0.95816 | -0.028081
|  |  | -8.5 | 0.10819 | 0.053055 | 0.0024233 | 0.059441 | 0.11062 | 0.10577 | -0.0063859 | 1 | 0.95619 | -0.05773 | 0.95793 | -0.030151
|  |  | -8 | 0.13252 | 0.065216 | 0.002873 | 0.072545 | 0.13539 | 0.12964 | -0.0073281 | 1 | 0.95756 | -0.054126 | 0.95909 | -0.028233
|  |  | -7.5 | 0.14923 | 0.074151 | 0.0031212 | 0.080224 | 0.15235 | 0.14611 | -0.006073 | 1 | 0.95903 | -0.039861 | 0.95986 | -0.02077
|  |  | -7 | 0.15852 | 0.079087 | 0.0030825 | 0.085282 | 0.16161 | 0.15544 | -0.006195 | 1 | 0.96185 | -0.038334 | 0.96262 | -0.019916
|  |  | -6.5 | 0.18168 | 0.090249 | 0.0031819 | 0.09916 | 0.18486 | 0.1785 | -0.0089111 | 1 | 0.96557 | -0.048204 | 0.96678 | -0.024941
|  |  | -6 | 0.23255 | 0.11384 | 0.0037975 | 0.12863 | 0.23634 | 0.22875 | -0.014794 | 1 | 0.96786 | -0.062596 | 0.96989 | -0.032292
|  |  | -5.5 | 0.25879 | 0.12685 | 0.0041857 | 0.14285 | 0.26298 | 0.25461 | -0.016007 | 1 | 0.96817 | -0.060869 | 0.97008 | -0.031394
|  |  | -5 | 0.29895 | 0.14857 | 0.004844 | 0.16785 | 0.3038 | 0.29411 | -0.01928 | 1 | 0.96811 | -0.063464 | 0.97019 | -0.03273
|  |  | -4.5 | 0.3664 | 0.18095 | 0.0057488 | 0.20798 | 0.37214 | 0.36065 | -0.027031 | 1 | 0.9691 | -0.072636 | 0.97182 | -0.037406
|  |  | -4 | 0.41889 | 0.20978 | 0.0059195 | 0.24066 | 0.42481 | 0.41297 | -0.030884 | 1 | 0.97213 | -0.072702 | 0.97485 | -0.037323
|  |  | -3.5 | 0.46344 | 0.23285 | 0.0067587 | 0.26505 | 0.4702 | 0.45668 | -0.032197 | 1 | 0.97125 | -0.068475 | 0.97366 | -0.035193
|  |  | -3 | 0.52521 | 0.26087 | 0.0073061 | 0.30908 | 0.53252 | 0.5179 | -0.048218 | 1 | 0.97256 | -0.090547 | 0.97677 | -0.046417
|  |  | -2.5 | 0.55573 | 0.2807 | 0.0071049 | 0.32498 | 0.56283 | 0.54862 | -0.044281 | 1 | 0.97475 | -0.078675 | 0.97792 | -0.040269
|  |  | -2 | 0.57483 | 0.28937 | 0.0069495 | 0.33585 | 0.58178 | 0.56788 | -0.046478 | 1 | 0.97611 | -0.079889 | 0.97937 | -0.040831
|  |  | -1.5 | 0.60614 | 0.30738 | 0.0066052 | 0.35413 | 0.61275 | 0.59954 | -0.046753 | 1 | 0.97844 | -0.0763 | 0.98141 | -0.038912
|  |  | -1 | 0.62055 | 0.31671 | 0.0064001 | 0.36218 | 0.62695 | 0.61415 | -0.045472 | 1 | 0.97958 | -0.072529 | 0.98226 | -0.036953
|  |  | -0.5 | 0.60089 | 0.30585 | 0.0060349 | 0.34979 | 0.60693 | 0.59486 | -0.043945 | 1 | 0.98011 | -0.072405 | 0.98278 | -0.03687
|  |  | 0 | 0.57593 | 0.29187 | 0.005352 | 0.33289 | 0.58128 | 0.57058 | -0.041016 | 1 | 0.98159 | -0.070561 | 0.98412 | -0.035881
|  |  | 0.5 | 0.54847 | 0.27851 | 0.0048018 | 0.31293 | 0.55327 | 0.54366 | -0.034424 | 1 | 0.98264 | -0.062219 | 0.98461 | -0.031617
|  |  | 1 | 0.49866 | 0.24628 | 0.0046616 | 0.28723 | 0.50332 | 0.494 | -0.040955 | 1 | 0.98148 | -0.08137 | 0.98484 | -0.041358
|  |  | 1.5 | 0.45099 | 0.23126 | 0.0034676 | 0.24942 | 0.45446 | 0.44752 | -0.018158 | 1 | 0.98474 | -0.039955 | 0.98555 | -0.020276
|  |  | 2 | 0.39105 | 0.19238 | 0.0036307 | 0.21909 | 0.39468 | 0.38742 | -0.026703 | 1 | 0.9816 | -0.067657 | 0.98393 | -0.034408
|  |  | 2.5 | 0.35059 | 0.1723 | 0.0029821 | 0.19559 | 0.35357 | 0.34761 | -0.023285 | 1 | 0.98313 | -0.065857 | 0.98533 | -0.033443
|  |  | 3 | 0.315 | 0.1568 | 0.002532 | 0.17157 | 0.31754 | 0.31247 | -0.014771 | 1 | 0.98405 | -0.046518 | 0.98515 | -0.023618
|  |  | 3.5 | 0.2666 | 0.13141 | 0.0022783 | 0.1445 | 0.26888 | 0.26432 | -0.013092 | 1 | 0.98305 | -0.048691 | 0.98426 | -0.024745
|  |  | 4 | 0.22992 | 0.11224 | 0.0020704 | 0.12479 | 0.23199 | 0.22785 | -0.012543 | 1 | 0.98215 | -0.054067 | 0.98364 | -0.027497
|  |  | 4.5 | 0.19934 | 0.097207 | 0.0018824 | 0.10666 | 0.20122 | 0.19746 | -0.0094524 | 1 | 0.98129 | -0.046974 | 0.98241 | -0.023917
|  |  | 5 | 0.17212 | 0.085064 | 0.0016468 | 0.090638 | 0.17377 | 0.17047 | -0.0055733 | 1 | 0.98105 | -0.032073 | 0.98157 | -0.016341
|  |  | 5.5 | 0.14246 | 0.069508 | 0.001627 | 0.075776 | 0.14408 | 0.14083 | -0.0062675 | 1 | 0.97742 | -0.043499 | 0.97838 | -0.022237
|  |  | 6 | 0.11725 | 0.057579 | 0.00141 | 0.061646 | 0.11866 | 0.11584 | -0.0040664 | 1 | 0.97623 | -0.03427 | 0.97684 | -0.017545
|  |  | 6.5 | 0.097813 | 0.04789 | 0.0010687 | 0.051056 | 0.098882 | 0.096744 | -0.0031662 | 1 | 0.97838 | -0.03202 | 0.97891 | -0.016358
|  |  | 7 | 0.077004 | 0.038063 | 0.0010493 | 0.03995 | 0.078053 | 0.075955 | -0.0018864 | 1 | 0.97311 | -0.024168 | 0.97341 | -0.012415
|  |  | 7.5 | 0.06004 | 0.030361 | 0.00090898 | 0.031059 | 0.060949 | 0.059131 | -0.0006981 | 1 | 0.97017 | -0.011454 | 0.97024 | -0.0059027
|  |  | 8 | 0.042397 | 0.022099 | 0.00078261 | 0.02108 | 0.043179 | 0.041614 | 0.0010186 | 1 | 0.96375 | 0.02359 | 0.96404 | 0.012236
|  |  | 8.5 | 0.024731 | 0.013081 | 0.00050366 | 0.011833 | 0.025234 | 0.024227 | 0.0012474 | 1 | 0.96008 | 0.049432 | 0.96135 | 0.025721
|  |  | 9 | 0.016926 | 0.0085698 | 0.00037742 | 0.0085717 | 0.017303 | 0.016548 | -1.9e-06 | 1 | 0.95638 | -0.00010981 | 0.95638 | -5.7407e-05
|  |  | 9.5 | 0.014534 | 0.0070629 | 0.00042057 | 0.0074768 | 0.014955 | 0.014114 | -0.0004139 | 1 | 0.94375 | -0.027677 | 0.94416 | -0.014659
|  |  | 10 | 0.0088196 | 0.0040503 | 0.00040817 | 0.0051728 | 0.0092278 | 0.0084115 | -0.0011225 | 1 | 0.91153 | -0.12164 | 0.91961 | -0.066331
|  |  | 10.5 | 0.0092235 | 0.0041523 | 0.00064707 | 0.0055485 | 0.0098706 | 0.0085764 | -0.0013962 | 1 | 0.86889 | -0.14145 | 0.88033 | -0.080688
|  |  | 11 | 0.0055213 | 0.0026484 | 0.00061417 | 0.0032959 | 0.0061355 | 0.0049072 | -0.00064754 | 1 | 0.7998 | -0.10554 | 0.80673 | -0.0656
|  |  | 11.5 | 0.0048514 | 0.002523 | 0.00059665 | 0.0026665 | 0.005448 | 0.0042547 | -0.00014353 | 1 | 0.78097 | -0.026345 | 0.78141 | -0.016861
|  |  | 12 | 0.0029054 | 0.0017219 | 0.000417 | 0.0015159 | 0.0033224 | 0.0024884 | 0.000206 | 1 | 0.74898 | 0.062004 | 0.75154 | 0.041298
|  |  | 12.5 | 0.0015731 | 0.000785 | 0.00024271 | 0.0010367 | 0.0018158 | 0.0013304 | -0.00025165 | 1 | 0.73267 | -0.13859 | 0.74566 | -0.093475
|  |  | 13 | 0.0010042 | 0.00061357 | 0.00030244 | 0.00069547 | 0.0013067 | 0.00070179 | -8.1897e-05 | 1 | 0.53709 | -0.062676 | 0.54073 | -0.058086
|  |  | 13.5 | 0.0010948 | 0.0010209 | 0.00056863 | 0.0006764 | 0.0016635 | 0.0005262 | 0.00034452 | 1 | 0.31633 | 0.20711 | 0.3781 | 0.28985
|  |  | 14 | 0.00088519 | 0.00067163 | 0.00036991 | 0.00055623 | 0.0012551 | 0.00051528 | 0.0001154 | 1 | 0.41055 | 0.091941 | 0.42072 | 0.11015
|  |  | 14.5 | 0.00088376 | 0.00071621 | 0.00034404 | 0.00050545 | 0.0012278 | 0.00053972 | 0.00021076 | 1 | 0.43958 | 0.17166 | 0.47191 | 0.18615
|  |  | 15 | 0.00080061 | 0.00064063 | 0.00030935 | 0.00043568 | 0.00111 | 0.00049126 | 0.00020495 | 1 | 0.44259 | 0.18465 | 0.47957 | 0.19762
|  |  | 15.5 | 0.00092757 | 0.00082922 | 0.0003357 | 0.00043339 | 0.0012633 | 0.00059188 | 0.00039584 | 1 | 0.46853 | 0.31334 | 0.56365 | 0.29473
|  |  | 16 | 0.00090343 | 0.00077272 | 0.00031185 | 0.00046504 | 0.0012153 | 0.00059158 | 0.00030768 | 1 | 0.48678 | 0.25318 | 0.54868 | 0.2398
|  |  | 16.5 | 0.0010188 | 0.00059963 | 0.00028706 | 0.00072837 | 0.0013058 | 0.00073171 | -0.00012875 | 1 | 0.56034 | -0.098593 | 0.56895 | -0.087084
|  |  | 17 | 0.00087762 | 0.00054646 | 0.00017163 | 0.00045407 | 0.0010492 | 0.000706 | 9.2389e-05 | 1 | 0.67286 | 0.088053 | 0.6786 | 0.065062
|  |  | 17.5 | 0.00082898 | 0.00044731 | 0.00014393 | 0.00049103 | 0.00097292 | 0.00068505 | -4.3722e-05 | 1 | 0.70412 | -0.044939 | 0.70556 | -0.031868
|  |  | 18 | 0.0010662 | 0.00053513 | 0.00015359 | 0.00071955 | 0.0012198 | 0.00091263 | -0.00018442 | 1 | 0.74817 | -0.15118 | 0.76329 | -0.099694
|  |  | 18.5 | 0.0010836 | 0.00055826 | 0.00016892 | 0.00066901 | 0.0012525 | 0.0009147 | -0.00011074 | 1 | 0.73028 | -0.088416 | 0.73561 | -0.060243
|  |  | 19 | 0.00082589 | 0.00051105 | 0.0001367 | 0.00043404 | 0.00096258 | 0.00068919 | 7.7011e-05 | 1 | 0.71598 | 0.080005 | 0.72043 | 0.05564
|  |  | 19.5 | 0.00085116 | 0.00040496 | 0.00011569 | 0.00053156 | 0.00096685 | 0.00073546 | -0.0001266 | 1 | 0.76068 | -0.13094 | 0.77187 | -0.085234
|  |  | 20 | 0.00076437 | 0.00043714 | 0.00010514 | 0.00041819 | 0.00086951 | 0.00065924 | 1.8954e-05 | 1 | 0.75817 | 0.021798 | 0.75849 | 0.014372
|  |  | 20.5 | 0.00073576 | 0.00039363 | 0.00011045 | 0.00042987 | 0.00084622 | 0.00062531 | -3.624e-05 | 1 | 0.73894 | -0.042826 | 0.74018 | -0.028945
|  |  | 21 | 0.0006373 | 0.00030113 | 0.00010097 | 0.00042284 | 0.00073827 | 0.00053633 | -0.00012171 | 1 | 0.72647 | -0.16486 | 0.74494 | -0.11158
|  |  | 21.5 | 0.00055576 | 0.00030709 | 8.8528e-05 | 0.00034428 | 0.00064429 | 0.00046723 | -3.7193e-05 | 1 | 0.72519 | -0.057728 | 0.72748 | -0.039718
|  |  | 22 | 0.00053597 | 0.00027383 | 8.183e-05 | 0.00033224 | 0.0006178 | 0.00045414 | -5.8412e-05 | 1 | 0.73509 | -0.094549 | 0.74115 | -0.06396
|  |  | 22.5 | 0.0004516 | 0.00024128 | 8.6628e-05 | 0.00029743 | 0.00053822 | 0.00036497 | -5.6148e-05 | 1 | 0.6781 | -0.10432 | 0.68607 | -0.076323
|  |  | 23 | 0.00044328 | 0.00023854 | 8.3596e-05 | 0.00027788 | 0.00052688 | 0.00035969 | -3.934e-05 | 1 | 0.68267 | -0.074666 | 0.68675 | -0.05447
|  |  | 23.5 | 0.00038389 | 0.00021076 | 7.8425e-05 | 0.00025308 | 0.00046231 | 0.00030546 | -4.232e-05 | 1 | 0.66073 | -0.09154 | 0.66704 | -0.068834
|  |  | 24 | 0.00037855 | 0.00021672 | 8.867e-05 | 0.00024462 | 0.00046722 | 0.00028988 | -2.7896e-05 | 1 | 0.62044 | -0.059706 | 0.6233 | -0.047969
|  |  | 24.5 | 0.00038925 | 0.00022149 | 8.9594e-05 | 0.00026703 | 0.00047884 | 0.00029966 | -4.5539e-05 | 1 | 0.62579 | -0.095102 | 0.63298 | -0.075408
|  |  | 25 | 0.00035784 | 0.00019533 | 8.3283e-05 | 0.00024164 | 0.00044112 | 0.00027456 | -4.6305e-05 | 1 | 0.6224 | -0.10497 | 0.63119 | -0.083541
|  |  | 25.5 | 0.00036192 | 0.00019179 | 8.5183e-05 | 0.00025058 | 0.0004471 | 0.00027674 | -5.8786e-05 | 1 | 0.61896 | -0.13148 | 0.63277 | -0.10466
|  |  | 26 | 0.00040004 | 0.000251 | 8.6763e-05 | 0.00024009 | 0.0004868 | 0.00031328 | 1.0909e-05 | 1 | 0.64354 | 0.02241 | 0.64393 | 0.017404
|  |  | 26.5 | 0.00039163 | 0.00022423 | 8.5339e-05 | 0.00024557 | 0.00047697 | 0.0003063 | -2.1339e-05 | 1 | 0.64216 | -0.044738 | 0.64372 | -0.034778
|  |  | 27 | 0.00037849 | 0.00020736 | 8.1346e-05 | 0.00023961 | 0.00045984 | 0.00029715 | -3.2247e-05 | 1 | 0.6462 | -0.070127 | 0.64999 | -0.05405
|  |  | 27.5 | 0.0003598 | 0.00022221 | 9.0875e-05 | 0.00023007 | 0.00045068 | 0.00026893 | -7.868e-06 | 1 | 0.59672 | -0.017458 | 0.59698 | -0.014624
|  |  | 28 | 0.00035182 | 0.00020415 | 8.0571e-05 | 0.00022948 | 0.00043239 | 0.00027125 | -2.5332e-05 | 1 | 0.62732 | -0.058586 | 0.63005 | -0.04656
|  |  | 28.5 | 0.00032255 | 0.00017685 | 8.8305e-05 | 0.00023198 | 0.00041086 | 0.00023425 | -5.5134e-05 | 1 | 0.57014 | -0.13419 | 0.58572 | -0.11558
|  |  | 29 | 0.00032291 | 0.00018853 | 7.6592e-05 | 0.00021076 | 0.0003995 | 0.00024632 | -2.2233e-05 | 1 | 0.61656 | -0.055652 | 0.61907 | -0.045009
|  |  | 29.5 | 0.00031543 | 0.00018793 | 8.1182e-05 | 0.00020969 | 0.00039661 | 0.00023425 | -2.1756e-05 | 1 | 0.59062 | -0.054855 | 0.59316 | -0.046305
|  |  | 30 | 0.00030297 | 0.00016964 | 7.4141e-05 | 0.00020886 | 0.00037711 | 0.00022883 | -3.922e-05 | 1 | 0.6068 | -0.104 | 0.61564 | -0.084872
|  |  | 30.5 | 0.00029871 | 0.00017458 | 7.5832e-05 | 0.00019547 | 0.00037454 | 0.00022288 | -2.0884e-05 | 1 | 0.59507 | -0.055759 | 0.59768 | -0.046714
|  |  | 31 | 0.00030333 | 0.00017274 | 8.6211e-05 | 0.00021565 | 0.00038954 | 0.00021712 | -4.2916e-05 | 1 | 0.55737 | -0.11017 | 0.56816 | -0.097573
|  |  | 31.5 | 0.00030193 | 0.00019437 | 8.9109e-05 | 0.00019378 | 0.00039104 | 0.00021282 | 5.96e-07 | 1 | 0.54424 | 0.0015241 | 0.54425 | 0.0014002
|  |  | 32 | 0.0002937 | 0.00018215 | 7.736e-05 | 0.00019169 | 0.00037106 | 0.00021634 | -9.536e-06 | 1 | 0.58304 | -0.025699 | 0.5836 | -0.022025
|  |  | 32.5 | 0.00030282 | 0.00017262 | 7.283e-05 | 0.00020278 | 0.00037565 | 0.00022999 | -3.016e-05 | 1 | 0.61225 | -0.080287 | 0.61749 | -0.065195
|  |  | 33 | 0.00027266 | 0.0001536 | 7.4626e-05 | 0.00018901 | 0.00034729 | 0.00019804 | -3.5405e-05 | 1 | 0.57024 | -0.10195 | 0.57928 | -0.088455
|  |  | 33.5 | 0.0002845 | 0.0001657 | 7.5132e-05 | 0.00019384 | 0.00035963 | 0.00020936 | -2.8133e-05 | 1 | 0.58217 | -0.078228 | 0.5874 | -0.066787
|  |  | 34 | 0.00027931 | 0.00016278 | 6.523e-05 | 0.00018841 | 0.00034454 | 0.00021408 | -2.563e-05 | 1 | 0.62135 | -0.074389 | 0.62579 | -0.059578
|  |  | 34.5 | 0.00026453 | 0.00015497 | 6.8002e-05 | 0.00018257 | 0.00033253 | 0.00019653 | -2.7597e-05 | 1 | 0.591 | -0.082991 | 0.5968 | -0.069756
|  |  | 35 | 0.00027317 | 0.00015992 | 7.0051e-05 | 0.00017995 | 0.00034322 | 0.00020312 | -2.0027e-05 | 1 | 0.5918 | -0.05835 | 0.59467 | -0.04914
|  |  | 35.5 | 0.00027442 | 0.0001564 | 7.172e-05 | 0.00018346 | 0.00034614 | 0.0002027 | -2.706e-05 | 1 | 0.5856 | -0.078176 | 0.5908 | -0.066356
|  |  | 36 | 0.00026348 | 0.00014228 | 5.9165e-05 | 0.00017953 | 0.00032265 | 0.00020432 | -3.7253e-05 | 1 | 0.63325 | -0.11546 | 0.64369 | -0.090173
|  |  | 36.5 | 0.00025213 | 0.00014019 | 6.1624e-05 | 0.0001716 | 0.00031375 | 0.00019051 | -3.1412e-05 | 1 | 0.60718 | -0.10012 | 0.61538 | -0.081708
|  |  | 37 | 0.00024578 | 0.00013691 | 6.0641e-05 | 0.00017023 | 0.00030642 | 0.00018514 | -3.3319e-05 | 1 | 0.6042 | -0.10874 | 0.61391 | -0.08903
|  |  | 37.5 | 0.00024098 | 0.00014222 | 6.0715e-05 | 0.00016338 | 0.0003017 | 0.00018027 | -2.1159e-05 | 1 | 0.59751 | -0.070133 | 0.60161 | -0.05842
|  |  | 38 | 0.00023213 | 0.00012904 | 6.1058e-05 | 0.00016528 | 0.00029319 | 0.00017107 | -3.624e-05 | 1 | 0.58349 | -0.12361 | 0.59644 | -0.10438
|  |  | 38.5 | 0.00023192 | 0.00012809 | 5.3488e-05 | 0.00015277 | 0.00028541 | 0.00017844 | -2.4677e-05 | 1 | 0.62519 | -0.086461 | 0.63114 | -0.068713
|  |  | 39 | 0.00021395 | 0.00011641 | 5.1201e-05 | 0.00014478 | 0.00026515 | 0.00016275 | -2.8373e-05 | 1 | 0.6138 | -0.10701 | 0.62306 | -0.0863
|  |  | 39.5 | 0.00020456 | 0.00011844 | 5.2974e-05 | 0.00013656 | 0.00025754 | 0.00015159 | -1.812e-05 | 1 | 0.58861 | -0.070359 | 0.5928 | -0.059484
|  |  | 40 | 0.00019193 | 0.00010347 | 4.7624e-05 | 0.00013512 | 0.00023955 | 0.0001443 | -3.1651e-05 | 1 | 0.60239 | -0.13213 | 0.61671 | -0.10796
|  |  | 40.5 | 0.00019479 | 0.00010747 | 4.8749e-05 | 0.00013441 | 0.00024354 | 0.00014604 | -2.6941e-05 | 1 | 0.59966 | -0.11062 | 0.60978 | -0.091213
|  |  | 41 | 0.00015873 | 9.4079e-05 | 4.22e-05 | 0.00010693 | 0.00020093 | 0.00011653 | -1.2852e-05 | 1 | 0.57995 | -0.063963 | 0.58347 | -0.054924
|  |  | 41.5 | 0.00010246 | 6.1755e-05 | 3.5241e-05 | 7.2763e-05 | 0.0001377 | 6.722e-05 | -1.1008e-05 | 1 | 0.48816 | -0.079941 | 0.49466 | -0.08116
|  |  | 42 | 7.0334e-05 | 4.5307e-05 | 2.6964e-05 | 5.1353e-05 | 9.7298e-05 | 4.337e-05 | -6.046e-06 | 1 | 0.44574 | -0.062139 | 0.45005 | -0.069256
|  |  | 42.5 | 5.284e-05 | 3.7588e-05 | 2.7299e-05 | 4.1913e-05 | 8.0139e-05 | 2.5541e-05 | -4.325e-06 | 1 | 0.31871 | -0.053969 | 0.32325 | -0.083872
|  |  | 43 | 4.1386e-05 | 3.1982e-05 | 2.297e-05 | 3.2701e-05 | 6.4356e-05 | 1.8416e-05 | -7.19e-07 | 1 | 0.28616 | -0.011172 | 0.28638 | -0.019511
|  |  | 43.5 | 3.3554e-05 | 3.3815e-05 | 3.4325e-05 | 3.2638e-05 | 6.7879e-05 | -7.71e-07 | 1.177e-06 | 1 | -0.011358 | 0.01734 | 0.020729 | -0.49544
|  |  | 44 | 2.9782e-05 | 2.3369e-05 | 2.5451e-05 | 3.1792e-05 | 5.5233e-05 | 4.331e-06 | -8.423e-06 | 1 | 0.078413 | -0.1525 | 0.17148 | -0.54793
|  |  | 44.5 | 3.0719e-05 | 2.2721e-05 | 2.2404e-05 | 2.6852e-05 | 5.3123e-05 | 8.315e-06 | -4.131e-06 | 1 | 0.15652 | -0.077763 | 0.17478 | -0.23055
|  |  | 45 | 2.932e-05 | 2.4952e-05 | 3.6761e-05 | 4.2636e-05 | 6.6081e-05 | -7.441e-06 | -1.7684e-05 | 1 | -0.1126 | -0.26761 | 0.29034 | 0.58625
|  |  | 45.5 | 2.2432e-05 | 1.7963e-05 | 1.7027e-05 | 2.0802e-05 | 3.9459e-05 | 5.405e-06 | -2.839e-06 | 1 | 0.13698 | -0.071948 | 0.15472 | -0.24182
|  |  | 46 | 2.308e-05 | 2.11e-05 | 2.0471e-05 | 2.3149e-05 | 4.3551e-05 | 2.609e-06 | -2.049e-06 | 1 | 0.059907 | -0.047048 | 0.076173 | -0.33287
|  |  | 46.5 | 2.3164e-05 | 2.3872e-05 | 2.6044e-05 | 2.5846e-05 | 4.9208e-05 | -2.88e-06 | -1.974e-06 | 1 | -0.058527 | -0.040115 | 0.070955 | 0.30044
|  |  | 47 | 1.9267e-05 | 1.8291e-05 | 1.6187e-05 | 1.9264e-05 | 3.5454e-05 | 3.08e-06 | -9.73e-07 | 1 | 0.086873 | -0.027444 | 0.091105 | -0.15299
|  |  | 47.5 | 1.9599e-05 | 1.964e-05 | 1.6317e-05 | 1.8977e-05 | 3.5916e-05 | 3.282e-06 | 6.63e-07 | 1 | 0.09138 | 0.01846 | 0.093226 | 0.099664
|  |  | 48 | 2.0059e-05 | 1.9036e-05 | 1.647e-05 | 1.562e-05 | 3.6529e-05 | 3.589e-06 | 3.416e-06 | 1 | 0.098251 | 0.093515 | 0.13564 | 0.38035
|  |  | 48.5 | 1.7792e-05 | 1.6339e-05 | 1.3661e-05 | 1.5032e-05 | 3.1453e-05 | 4.131e-06 | 1.307e-06 | 1 | 0.13134 | 0.041554 | 0.13776 | 0.15321
|  |  | 49 | 1.7524e-05 | 1.8131e-05 | 1.7289e-05 | 1.4957e-05 | 3.4813e-05 | 2.35e-07 | 3.174e-06 | 1 | 0.0067504 | 0.091173 | 0.091422 | 0.74845
|  |  | 49.5 | 1.7967e-05 | 1.7051e-05 | 1.8101e-05 | 2.0474e-05 | 3.6068e-05 | -1.34e-07 | -3.423e-06 | 1 | -0.0037152 | -0.094904 | 0.094977 | 0.76583
|  |  | 50 | 1.9264e-05 | 1.6518e-05 | 2.2091e-05 | 2.2706e-05 | 4.1355e-05 | -2.827e-06 | -6.188e-06 | 1 | -0.068359 | -0.14963 | 0.16451 | 0.57113
|  |  | 50.5 | 1.673e-05 | 1.6194e-05 | 1.5315e-05 | 1.4912e-05 | 3.2045e-05 | 1.415e-06 | 1.282e-06 | 1 | 0.044157 | 0.040006 | 0.059584 | 0.36806
|  |  | 51 | 1.9148e-05 | 2.0277e-05 | 1.7036e-05 | 1.6924e-05 | 3.6184e-05 | 2.112e-06 | 3.353e-06 | 1 | 0.058368 | 0.092665 | 0.10952 | 0.50435
|  |  | 51.5 | 1.7928e-05 | 1.4175e-05 | 1.2644e-05 | 1.6213e-05 | 3.0572e-05 | 5.284e-06 | -2.038e-06 | 1 | 0.17284 | -0.066662 | 0.18525 | -0.18406
|  |  | 52 | 1.5074e-05 | 1.3478e-05 | 1.3456e-05 | 1.511e-05 | 2.853e-05 | 1.618e-06 | -1.632e-06 | 1 | 0.056712 | -0.057203 | 0.080551 | -0.39485
|  |  | 52.5 | 1.7021e-05 | 1.5274e-05 | 1.6034e-05 | 1.6704e-05 | 3.3055e-05 | 9.87e-07 | -1.43e-06 | 1 | 0.029859 | -0.043261 | 0.052565 | -0.48334
|  |  | 53 | 1.6259e-05 | 1.5658e-05 | 1.4462e-05 | 1.4719e-05 | 3.0721e-05 | 1.797e-06 | 9.39e-07 | 1 | 0.058494 | 0.030565 | 0.065999 | 0.24076
|  |  | 53.5 | 1.4519e-05 | 1.5143e-05 | 1.5728e-05 | 1.5915e-05 | 3.0247e-05 | -1.209e-06 | -7.72e-07 | 1 | -0.039971 | -0.025523 | 0.047425 | 0.28414
|  |  | 54 | 1.4914e-05 | 1.6727e-05 | 1.6291e-05 | 1.4599e-05 | 3.1205e-05 | -1.377e-06 | 2.128e-06 | 1 | -0.044128 | 0.068194 | 0.081226 | -0.49824
|  |  | 54.5 | 1.853e-05 | 1.546e-05 | 1.5918e-05 | 1.7472e-05 | 3.4448e-05 | 2.612e-06 | -2.012e-06 | 1 | 0.075824 | -0.058407 | 0.095712 | -0.32818
|  |  | 55 | 1.8811e-05 | 1.7528e-05 | 1.5658e-05 | 1.6872e-05 | 3.4469e-05 | 3.153e-06 | 6.56e-07 | 1 | 0.091473 | 0.019032 | 0.093432 | 0.10256
|  |  | 55.5 | 1.6963e-05 | 1.4115e-05 | 1.2349e-05 | 1.3687e-05 | 2.9312e-05 | 4.614e-06 | 4.28e-07 | 1 | 0.15741 | 0.014602 | 0.15809 | 0.046248
|  |  | 56 | 1.7628e-05 | 1.5426e-05 | 1.1131e-05 | 1.3988e-05 | 2.8759e-05 | 6.497e-06 | 1.438e-06 | 1 | 0.22591 | 0.050002 | 0.23138 | 0.10891
|  |  | 56.5 | 1.5838e-05 | 1.527e-05 | 1.2722e-05 | 1.261e-05 | 2.856e-05 | 3.116e-06 | 2.66e-06 | 1 | 0.1091 | 0.093137 | 0.14345 | 0.35331
|  |  | 57 | 1.8429e-05 | 1.6913e-05 | 1.5684e-05 | 1.7896e-05 | 3.4113e-05 | 2.745e-06 | -9.83e-07 | 1 | 0.080468 | -0.028816 | 0.085472 | -0.17194
|  |  | 57.5 | 1.6695e-05 | 1.4208e-05 | 1.489e-05 | 1.7025e-05 | 3.1585e-05 | 1.805e-06 | -2.817e-06 | 1 | 0.057147 | -0.089188 | 0.10593 | -0.50047
|  |  | 58 | 1.4583e-05 | 1.3873e-05 | 1.1541e-05 | 1.2536e-05 | 2.6124e-05 | 3.042e-06 | 1.337e-06 | 1 | 0.11644 | 0.051179 | 0.1272 | 0.20705
|  |  | 58.5 | 1.3666e-05 | 1.1996e-05 | 1.4227e-05 | 1.5482e-05 | 2.7893e-05 | -5.61e-07 | -3.486e-06 | 1 | -0.020113 | -0.12498 | 0.12659 | 0.70562
|  |  | 59 | 1.3832e-05 | 1.4e-05 | 1.1843e-05 | 1.1459e-05 | 2.5675e-05 | 1.989e-06 | 2.541e-06 | 1 | 0.077468 | 0.098968 | 0.12568 | 0.45333
|  |  | 59.5 | 1.1442e-05 | 1.1303e-05 | 1.2454e-05 | 1.2312e-05 | 2.3896e-05 | -1.012e-06 | -1.009e-06 | 1 | -0.04235 | -0.042225 | 0.059803 | 0.39196
|  |  | 60 | 9.451e-06 | 8.867e-06 | 8.532e-06 | 8.946e-06 | 1.7983e-05 | 9.19e-07 | -7.9e-08 | 1 | 0.051104 | -0.004393 | 0.051292 | -0.042876
Note on A13: FO: The setup to measure the optical power output vs the angle is composed by a motorized rotation stage (central hub) with an arm. The detection sits at the end of the arm and is composed by a Glan-Thompson polarizer mounted in a motorized rotation stage. An aperture with 6.5 mm is monted after the polarized to reject the undesired polarization. It is followed by a slit with a 1 mm width and a detector. The distance between central hub and the slit is ~573mm (1 mm slit = 0,1º). The detector is connected to a lock-in amplifer SR830. The laser is installed in the diode holder with the laser diode window centred with the central hub. A wave generator is used to drive the lase diode controller with a sinosoidal wave of amplitude 'A' and offset of 'A/2'. The power meter was used to adjust the average power. The laser was measured with the diode die oriented roughly with the horizontal and vertical of the laboratory to measure the angular dispertion in both orientations. Polariation was measured at 0º (horizontal), 45º, 90º and 135º.

## Angle_Perpendicular
Optical Power Output vs Angle |  |  | Perpendicular Direction
Measurment Information: |  |  | Measured Data: |  |  |  |  | Calculated Data
|  |  | PSA = Polarization State Analyser |  |  |  |  | * DOLP - degree of linear polarization
PowerMeter: | Thorlabs PM320E
PowerMeter Detector: | Thorlabs S142C |  | Angle | Intensity | Intensity | Intensity | Intensity | Stocks Vector Components |  |  | Norm. Stockes Vector Comp. |  |  | DOLP* | Azimuth
Laser Diode Controller: | Thorlabs LDC205C |  | (⁰) | (a.u.) | (a.u.) | (a.u.) | (a.u.) | s0 | s1 | s2 | S0 | S1 | S2 |  | (⁰)
Wave Generator: | Teledyne LeCroy WaveStation 2023 |  |  | PSA = 0⁰ | PSA = 45⁰ | PSA = 90⁰ | PSA = 135⁰
Temperature Controller: | Thorlabs TED200C |  | -60 | 7.5263e-05 | 2.2229e-05 | 6.4124e-05 | 0.00010369 | 0.00013939 | 1.1139e-05 | -8.1461e-05 | 1 | 0.079914 | -0.58442 | 0.58986 | -0.71745
Laser Diode Holder: | Thorlabs TCLDM9 |  | -59.5 | 0.000108 | 2.2907e-05 | 8.2404e-05 | 0.00016107 | 0.00019041 | 2.56e-05 | -0.00013817 | 1 | 0.13445 | -0.72564 | 0.73799 | -0.6938
Detector | Hamamatsu H9656-20 |  | -59 | 0.00020021 | 2.5578e-05 | 0.00018862 | 0.00035602 | 0.00038883 | 1.1593e-05 | -0.00033044 | 1 | 0.029815 | -0.84983 | 0.85035 | -0.76786
Lock-in Amplifier | Stanford Research System SR830 |  | -58.5 | 0.00020009 | 1.7876e-05 | 0.00016093 | 0.00034392 | 0.00036103 | 3.9161e-05 | -0.00032605 | 1 | 0.10847 | -0.9031 | 0.9096 | -0.72563
Measurment Notes |  |  | -58 | 0.00013 | 2.1674e-05 | 0.00010359 | 0.00020325 | 0.00023359 | 2.6406e-05 | -0.00018158 | 1 | 0.11304 | -0.77733 | 0.78551 | -0.71319
|  |  | -57.5 | 0.00010565 | 2.4842e-05 | 8.8037e-05 | 0.00016487 | 0.00019369 | 1.7613e-05 | -0.00014003 | 1 | 0.090935 | -0.72294 | 0.72864 | -0.72283
Fixed LD house Temp.: | 20 | ⁰C | -57 | 0.00016856 | 2.0702e-05 | 0.00015351 | 0.0002991 | 0.00032208 | 1.505e-05 | -0.0002784 | 1 | 0.046728 | -0.86438 | 0.86564 | -0.75839
Optical Power (avg.): | 5.2505 | mW | -56.5 | 0.00021818 | 2.5082e-05 | 0.00014928 | 0.00031716 | 0.00036746 | 6.8903e-05 | -0.00029208 | 1 | 0.18751 | -0.79484 | 0.81666 | -0.66956
LD Photodiode Current: | 0.016 | mA | -56 | 0.00029266 | 2.9761e-05 | 0.0001685 | 0.00044197 | 0.00046116 | 0.00012416 | -0.00041221 | 1 | 0.26923 | -0.89384 | 0.93351 | -0.63912
Modulation Freq.: | 275 | Hz | -55.5 | 0.00022259 | 2.6995e-05 | 0.0001773 | 0.00040072 | 0.00039989 | 4.5299e-05 | -0.00037373 | 1 | 0.11328 | -0.93458 | 0.94142 | -0.72509
|  |  | -55 | 0.00018334 | 2.3879e-05 | 0.000128 | 0.00030387 | 0.00031135 | 5.5343e-05 | -0.00027999 | 1 | 0.17775 | -0.89928 | 0.91668 | -0.68782
FWHM | 31.8919 | ⁰ | -54.5 | 0.0004524 | 2.7673e-05 | 0.00028226 | 0.0006814 | 0.00073466 | 0.00017014 | -0.00065373 | 1 | 0.23159 | -0.88984 | 0.91948 | -0.65809
DOLP (max value) | 98.5597 | % | -54 | 0.00045637 | 3.4265e-05 | 0.00030705 | 0.00076169 | 0.00076342 | 0.00014931 | -0.00072743 | 1 | 0.19558 | -0.95285 | 0.97272 | -0.68418
|  |  | -53.5 | 0.00016388 | 2.0147e-05 | 0.00012756 | 0.00027466 | 0.00029144 | 3.6329e-05 | -0.00025451 | 1 | 0.12465 | -0.8733 | 0.88215 | -0.71451
Auxiliar |  |  | -53 | 9.5815e-05 | 2.2104e-05 | 0.00010505 | 0.00018805 | 0.00020087 | -9.239e-06 | -0.00016595 | 1 | -0.045995 | -0.82616 | 0.82744 | 0.75759
Max. Intentity PSA = 0º | 0.36621 |  | -52.5 | 0.00013054 | 2.2946e-05 | 0.00013539 | 0.00023896 | 0.00026593 | -4.858e-06 | -0.00021601 | 1 | -0.018268 | -0.81229 | 0.81249 | 0.77416
Row Position = Max. | 128 |  | -52 | 0.00018048 | 2.1532e-05 | 0.00013956 | 0.00031448 | 0.00032005 | 4.0919e-05 | -0.00029294 | 1 | 0.12785 | -0.91531 | 0.9242 | -0.71601
Half-Maximum | 0.18311 |  | -51.5 | 0.00015175 | 2.4406e-05 | 0.00010568 | 0.00024051 | 0.00025743 | 4.6074e-05 | -0.0002161 | 1 | 0.17897 | -0.83944 | 0.85831 | -0.68037
Negative 'Half-angle' | -17.223 | ⁰ | -51 | 0.00010243 | 2.5125e-05 | 9.6262e-05 | 0.00017458 | 0.00019869 | 6.169e-06 | -0.00014946 | 1 | 0.031048 | -0.75221 | 0.75285 | -0.76477
Positive 'Half-angle' | 14.6689 | ⁰ | -50.5 | 0.0001536 | 2.7497e-05 | 0.00012952 | 0.00026912 | 0.00028312 | 2.408e-05 | -0.00024162 | 1 | 0.085051 | -0.85341 | 0.85763 | -0.73573
|  |  | -50 | 0.00014526 | 2.7446e-05 | 0.00013325 | 0.00025326 | 0.0002785 | 1.201e-05 | -0.00022582 | 1 | 0.043123 | -0.81082 | 0.81196 | -0.75883
|  |  | -49.5 | 0.00011909 | 2.5911e-05 | 9.206e-05 | 0.00019258 | 0.00021115 | 2.7031e-05 | -0.00016667 | 1 | 0.12802 | -0.78935 | 0.79967 | -0.70501
|  |  | -49 | 0.00011882 | 2.2406e-05 | 0.00012261 | 0.00021797 | 0.00024143 | -3.786e-06 | -0.00019557 | 1 | -0.015682 | -0.81004 | 0.8102 | 0.77572
|  |  | -48.5 | 0.00020626 | 3.1796e-05 | 0.00015453 | 0.000314 | 0.00036079 | 5.1737e-05 | -0.0002822 | 1 | 0.1434 | -0.78218 | 0.79522 | -0.69474
|  |  | -48 | 0.00027293 | 3.4865e-05 | 0.00020981 | 0.00044942 | 0.00048274 | 6.3123e-05 | -0.00041456 | 1 | 0.13076 | -0.85875 | 0.86865 | -0.70985
|  |  | -47.5 | 0.000173 | 3.5342e-05 | 0.00014868 | 0.00028372 | 0.00032169 | 2.4318e-05 | -0.00024838 | 1 | 0.075595 | -0.77211 | 0.7758 | -0.7366
|  |  | -47 | 0.00019545 | 3.2919e-05 | 0.0001868 | 0.00035703 | 0.00038225 | 8.643e-06 | -0.00032411 | 1 | 0.022611 | -0.84792 | 0.84822 | -0.77207
|  |  | -46.5 | 0.00035367 | 3.4355e-05 | 0.00031382 | 0.00063736 | 0.00066749 | 3.9846e-05 | -0.000603 | 1 | 0.059696 | -0.90339 | 0.90536 | -0.75241
|  |  | -46 | 0.00049973 | 4.8489e-05 | 0.0004566 | 0.00096518 | 0.00095633 | 4.3126e-05 | -0.00091669 | 1 | 0.045095 | -0.95855 | 0.95961 | -0.76189
|  |  | -45.5 | 0.00074864 | 7.2323e-05 | 0.00071299 | 0.0013611 | 0.0014616 | 3.5644e-05 | -0.0012888 | 1 | 0.024386 | -0.88177 | 0.8821 | -0.77157
|  |  | -45 | 0.0013986 | 0.0001257 | 0.0012617 | 0.0026441 | 0.0026603 | 0.00013685 | -0.0025184 | 1 | 0.051442 | -0.94666 | 0.94805 | -0.75825
|  |  | -44.5 | 0.0025768 | 0.00023115 | 0.00242 | 0.0047551 | 0.0049968 | 0.00015688 | -0.0045239 | 1 | 0.031396 | -0.90536 | 0.9059 | -0.76807
|  |  | -44 | 0.0052157 | 0.00034279 | 0.0045905 | 0.009801 | 0.0098062 | 0.00062513 | -0.0094582 | 1 | 0.063748 | -0.96451 | 0.96661 | -0.7524
|  |  | -43.5 | 0.0112 | 0.00038272 | 0.009295 | 0.020458 | 0.020495 | 0.001905 | -0.020076 | 1 | 0.092947 | -0.97953 | 0.98393 | -0.7381
|  |  | -43 | 0.020662 | 0.00043148 | 0.017504 | 0.037735 | 0.038166 | 0.0031585 | -0.037304 | 1 | 0.082756 | -0.9774 | 0.9809 | -0.74316
|  |  | -42.5 | 0.026343 | 0.00049109 | 0.021504 | 0.046269 | 0.047846 | 0.0048389 | -0.045778 | 1 | 0.10113 | -0.95677 | 0.9621 | -0.73274
|  |  | -42 | 0.019308 | 0.00046468 | 0.016087 | 0.034462 | 0.035395 | 0.0032215 | -0.033998 | 1 | 0.091016 | -0.96052 | 0.96482 | -0.73816
|  |  | -41.5 | 0.022755 | 0.00051284 | 0.018721 | 0.041088 | 0.041475 | 0.0040341 | -0.040575 | 1 | 0.097265 | -0.9783 | 0.98312 | -0.73585
|  |  | -41 | 0.023304 | 0.00055212 | 0.018696 | 0.040959 | 0.042 | 0.0046081 | -0.040406 | 1 | 0.10972 | -0.96206 | 0.96829 | -0.72862
|  |  | -40.5 | 0.02689 | 0.00058282 | 0.021931 | 0.048226 | 0.048821 | 0.0049591 | -0.047643 | 1 | 0.10158 | -0.97587 | 0.98114 | -0.73354
|  |  | -40 | 0.022856 | 0.0005712 | 0.018683 | 0.040898 | 0.041538 | 0.0041733 | -0.040326 | 1 | 0.10047 | -0.97082 | 0.976 | -0.73384
|  |  | -39.5 | 0.026522 | 0.00062257 | 0.021408 | 0.046528 | 0.04793 | 0.0051136 | -0.045906 | 1 | 0.10669 | -0.95776 | 0.96369 | -0.72993
|  |  | -39 | 0.02968 | 0.00066156 | 0.023806 | 0.051937 | 0.053486 | 0.0058747 | -0.051276 | 1 | 0.10984 | -0.95867 | 0.96495 | -0.72836
|  |  | -38.5 | 0.032582 | 0.00070185 | 0.026201 | 0.057179 | 0.058783 | 0.0063802 | -0.056477 | 1 | 0.10854 | -0.96077 | 0.96688 | -0.72915
|  |  | -38 | 0.033696 | 0.00072724 | 0.027432 | 0.059632 | 0.061127 | 0.0062639 | -0.058904 | 1 | 0.10247 | -0.96364 | 0.96907 | -0.73243
|  |  | -37.5 | 0.034144 | 0.00072903 | 0.027928 | 0.06017 | 0.062071 | 0.0062162 | -0.05944 | 1 | 0.10015 | -0.95762 | 0.96284 | -0.7333
|  |  | -37 | 0.032219 | 0.00075847 | 0.026341 | 0.057198 | 0.05856 | 0.0058786 | -0.056439 | 1 | 0.10039 | -0.96379 | 0.969 | -0.73351
|  |  | -36.5 | 0.035332 | 0.00079924 | 0.028849 | 0.062794 | 0.064181 | 0.0064832 | -0.061995 | 1 | 0.10101 | -0.96594 | 0.97121 | -0.7333
|  |  | -36 | 0.03792 | 0.00083834 | 0.031281 | 0.067547 | 0.069201 | 0.0066394 | -0.066709 | 1 | 0.095944 | -0.96399 | 0.96875 | -0.7358
|  |  | -35.5 | 0.03629 | 0.00083626 | 0.030434 | 0.065568 | 0.066723 | 0.0058557 | -0.064731 | 1 | 0.087761 | -0.97014 | 0.97411 | -0.74029
|  |  | -35 | 0.042671 | 0.00087786 | 0.035378 | 0.075596 | 0.078049 | 0.0072937 | -0.074718 | 1 | 0.09345 | -0.95733 | 0.96188 | -0.73674
|  |  | -34.5 | 0.040594 | 0.0008986 | 0.034191 | 0.07367 | 0.074786 | 0.0064029 | -0.072771 | 1 | 0.085617 | -0.97306 | 0.97682 | -0.74152
|  |  | -34 | 0.041432 | 0.00096673 | 0.034102 | 0.073147 | 0.075533 | 0.00733 | -0.072181 | 1 | 0.097043 | -0.95561 | 0.96053 | -0.7348
|  |  | -33.5 | 0.046234 | 0.00096143 | 0.039072 | 0.082497 | 0.085307 | 0.0071621 | -0.081536 | 1 | 0.083957 | -0.95579 | 0.95947 | -0.74159
|  |  | -33 | 0.050141 | 0.0010626 | 0.041722 | 0.089135 | 0.091862 | 0.008419 | -0.088072 | 1 | 0.091648 | -0.95874 | 0.96311 | -0.73775
|  |  | -32.5 | 0.052048 | 0.0011401 | 0.042576 | 0.091031 | 0.094624 | 0.0094719 | -0.08989 | 1 | 0.1001 | -0.94997 | 0.95523 | -0.73291
|  |  | -32 | 0.052864 | 0.0011461 | 0.044073 | 0.093773 | 0.096938 | 0.008791 | -0.092627 | 1 | 0.090687 | -0.95554 | 0.95983 | -0.73809
|  |  | -31.5 | 0.053551 | 0.0012276 | 0.044213 | 0.095437 | 0.097764 | 0.0093384 | -0.094209 | 1 | 0.09552 | -0.96364 | 0.96836 | -0.736
|  |  | -31 | 0.057869 | 0.0012739 | 0.04789 | 0.10349 | 0.10576 | 0.0099793 | -0.10221 | 1 | 0.094359 | -0.96646 | 0.97106 | -0.73674
|  |  | -30.5 | 0.064774 | 0.0013254 | 0.053864 | 0.11392 | 0.11864 | 0.01091 | -0.1126 | 1 | 0.091962 | -0.94909 | 0.95353 | -0.7371
|  |  | -30 | 0.060784 | 0.0013866 | 0.050812 | 0.10904 | 0.1116 | 0.0099716 | -0.10765 | 1 | 0.089355 | -0.96467 | 0.9688 | -0.73922
|  |  | -29.5 | 0.065072 | 0.0014606 | 0.053444 | 0.11411 | 0.11852 | 0.011627 | -0.11265 | 1 | 0.098109 | -0.95047 | 0.95552 | -0.73397
|  |  | -29 | 0.072098 | 0.0015268 | 0.06041 | 0.12891 | 0.13251 | 0.011689 | -0.12738 | 1 | 0.08821 | -0.9613 | 0.96534 | -0.73965
|  |  | -28.5 | 0.072564 | 0.0016255 | 0.059891 | 0.12769 | 0.13245 | 0.012673 | -0.12606 | 1 | 0.095675 | -0.95173 | 0.95653 | -0.7353
|  |  | -28 | 0.076546 | 0.0016847 | 0.062935 | 0.13373 | 0.13948 | 0.013611 | -0.13204 | 1 | 0.097581 | -0.94668 | 0.9517 | -0.73404
|  |  | -27.5 | 0.081116 | 0.0017524 | 0.067704 | 0.14371 | 0.14882 | 0.013412 | -0.14196 | 1 | 0.090126 | -0.95388 | 0.95812 | -0.7383
|  |  | -27 | 0.082192 | 0.0018387 | 0.069069 | 0.14667 | 0.15126 | 0.013122 | -0.14483 | 1 | 0.086754 | -0.95748 | 0.9614 | -0.74022
|  |  | -26.5 | 0.091347 | 0.0019555 | 0.076058 | 0.16001 | 0.16741 | 0.015289 | -0.15805 | 1 | 0.091331 | -0.94411 | 0.94852 | -0.73718
|  |  | -26 | 0.087037 | 0.0020008 | 0.073105 | 0.15509 | 0.16014 | 0.013931 | -0.15309 | 1 | 0.086993 | -0.95597 | 0.95992 | -0.74002
|  |  | -25.5 | 0.09163 | 0.0020609 | 0.077546 | 0.16361 | 0.16918 | 0.014084 | -0.16155 | 1 | 0.08325 | -0.9549 | 0.95852 | -0.74192
|  |  | -25 | 0.099732 | 0.0021665 | 0.084763 | 0.17868 | 0.18449 | 0.014969 | -0.17651 | 1 | 0.081134 | -0.95674 | 0.96018 | -0.7431
|  |  | -24.5 | 0.10733 | 0.0022695 | 0.091164 | 0.19 | 0.1985 | 0.016167 | -0.18773 | 1 | 0.081447 | -0.94578 | 0.94928 | -0.74245
|  |  | -24 | 0.11502 | 0.0024049 | 0.097474 | 0.20477 | 0.21249 | 0.017547 | -0.20237 | 1 | 0.082578 | -0.95235 | 0.95592 | -0.74215
|  |  | -23.5 | 0.11551 | 0.0025139 | 0.096962 | 0.20319 | 0.21247 | 0.018548 | -0.20067 | 1 | 0.087294 | -0.94447 | 0.94849 | -0.73932
|  |  | -23 | 0.1175 | 0.0026181 | 0.099282 | 0.2078 | 0.21678 | 0.018219 | -0.20518 | 1 | 0.084044 | -0.94646 | 0.95019 | -0.74112
|  |  | -22.5 | 0.11982 | 0.0027564 | 0.10139 | 0.21289 | 0.22121 | 0.018432 | -0.21014 | 1 | 0.083324 | -0.94995 | 0.95359 | -0.74165
|  |  | -22 | 0.12882 | 0.0028782 | 0.10899 | 0.22791 | 0.23781 | 0.019822 | -0.22503 | 1 | 0.083352 | -0.94625 | 0.94991 | -0.74147
|  |  | -21.5 | 0.1378 | 0.0029979 | 0.11703 | 0.24494 | 0.25482 | 0.020767 | -0.24194 | 1 | 0.081496 | -0.94943 | 0.95292 | -0.74258
|  |  | -21 | 0.14217 | 0.0031667 | 0.12066 | 0.25247 | 0.26283 | 0.021516 | -0.24931 | 1 | 0.081862 | -0.94854 | 0.95206 | -0.74235
|  |  | -20.5 | 0.14765 | 0.0032547 | 0.12454 | 0.25931 | 0.27219 | 0.023118 | -0.25606 | 1 | 0.084934 | -0.94073 | 0.94456 | -0.74038
|  |  | -20 | 0.157 | 0.0034919 | 0.13188 | 0.27365 | 0.28888 | 0.025116 | -0.27016 | 1 | 0.086942 | -0.9352 | 0.93923 | -0.73905
|  |  | -19.5 | 0.16273 | 0.0035687 | 0.13667 | 0.28333 | 0.29939 | 0.026062 | -0.27976 | 1 | 0.087049 | -0.93442 | 0.93846 | -0.73895
|  |  | -19 | 0.16499 | 0.0037677 | 0.13866 | 0.28913 | 0.30364 | 0.026329 | -0.28536 | 1 | 0.08671 | -0.93977 | 0.94376 | -0.7394
|  |  | -18.5 | 0.17375 | 0.0039158 | 0.14563 | 0.3017 | 0.31938 | 0.028122 | -0.29778 | 1 | 0.088051 | -0.93237 | 0.93652 | -0.73832
|  |  | -18 | 0.16917 | 0.0041452 | 0.14149 | 0.29389 | 0.31066 | 0.027687 | -0.28974 | 1 | 0.089122 | -0.93265 | 0.9369 | -0.73776
|  |  | -17.5 | 0.19082 | 0.0041991 | 0.16077 | 0.33222 | 0.35159 | 0.030052 | -0.32802 | 1 | 0.085475 | -0.93296 | 0.93687 | -0.73972
|  |  | -17 | 0.1871 | 0.004292 | 0.15717 | 0.32388 | 0.34427 | 0.029938 | -0.31959 | 1 | 0.086961 | -0.92832 | 0.93238 | -0.7387
|  |  | -16.5 | 0.19647 | 0.0045498 | 0.16709 | 0.3429 | 0.36357 | 0.029381 | -0.33835 | 1 | 0.080814 | -0.93064 | 0.93414 | -0.74209
|  |  | -16 | 0.20868 | 0.0046864 | 0.17554 | 0.35971 | 0.38422 | 0.033142 | -0.35503 | 1 | 0.086258 | -0.92402 | 0.92804 | -0.73886
|  |  | -15.5 | 0.20593 | 0.0047965 | 0.17416 | 0.35599 | 0.38009 | 0.031777 | -0.35119 | 1 | 0.083604 | -0.92397 | 0.92774 | -0.74028
|  |  | -15 | 0.20978 | 0.0048833 | 0.17912 | 0.36536 | 0.3889 | 0.030655 | -0.36047 | 1 | 0.078824 | -0.9269 | 0.93025 | -0.74298
|  |  | -14.5 | 0.22791 | 0.0051613 | 0.19434 | 0.39478 | 0.42224 | 0.033569 | -0.38962 | 1 | 0.079502 | -0.92273 | 0.92615 | -0.74242
|  |  | -14 | 0.24072 | 0.0054083 | 0.20508 | 0.41596 | 0.4458 | 0.035645 | -0.41055 | 1 | 0.079957 | -0.92092 | 0.92438 | -0.7421
|  |  | -13.5 | 0.24457 | 0.0055294 | 0.20856 | 0.42291 | 0.45313 | 0.036011 | -0.41739 | 1 | 0.079472 | -0.92112 | 0.92454 | -0.74237
|  |  | -13 | 0.25488 | 0.0057459 | 0.21558 | 0.43305 | 0.47046 | 0.039308 | -0.4273 | 1 | 0.083552 | -0.90826 | 0.91209 | -0.73953
|  |  | -12.5 | 0.2619 | 0.0059052 | 0.2229 | 0.45044 | 0.48481 | 0.039003 | -0.44454 | 1 | 0.080451 | -0.91694 | 0.92046 | -0.74164
|  |  | -12 | 0.26911 | 0.0061102 | 0.22894 | 0.45929 | 0.49805 | 0.040162 | -0.45318 | 1 | 0.080638 | -0.90991 | 0.91348 | -0.7412
|  |  | -11.5 | 0.26636 | 0.0062161 | 0.2273 | 0.45746 | 0.49366 | 0.039063 | -0.45124 | 1 | 0.07913 | -0.91409 | 0.91751 | -0.74222
|  |  | -11 | 0.27301 | 0.0063238 | 0.23358 | 0.46735 | 0.50659 | 0.03943 | -0.46102 | 1 | 0.077834 | -0.91005 | 0.91337 | -0.74274
|  |  | -10.5 | 0.2909 | 0.0066309 | 0.24695 | 0.49265 | 0.53784 | 0.043946 | -0.48602 | 1 | 0.081708 | -0.90364 | 0.90732 | -0.74031
|  |  | -10 | 0.28577 | 0.0068188 | 0.24396 | 0.49179 | 0.52973 | 0.041811 | -0.48497 | 1 | 0.078929 | -0.91552 | 0.91891 | -0.7424
|  |  | -9.5 | 0.29724 | 0.007 | 0.2525 | 0.50428 | 0.54975 | 0.044739 | -0.49728 | 1 | 0.081381 | -0.90455 | 0.90821 | -0.74053
|  |  | -9 | 0.30915 | 0.0071345 | 0.26264 | 0.52112 | 0.57178 | 0.046509 | -0.51399 | 1 | 0.081341 | -0.89892 | 0.9026 | -0.74028
|  |  | -8.5 | 0.30902 | 0.0072708 | 0.26282 | 0.52637 | 0.57184 | 0.046204 | -0.5191 | 1 | 0.080799 | -0.90777 | 0.91136 | -0.74101
|  |  | -8 | 0.30683 | 0.0073548 | 0.26233 | 0.52533 | 0.56916 | 0.044495 | -0.51798 | 1 | 0.078177 | -0.91008 | 0.91343 | -0.74255
|  |  | -7.5 | 0.31067 | 0.0075064 | 0.26563 | 0.53064 | 0.5763 | 0.045044 | -0.52314 | 1 | 0.078161 | -0.90775 | 0.91111 | -0.74245
|  |  | -7 | 0.32288 | 0.0075541 | 0.27741 | 0.55201 | 0.60028 | 0.045471 | -0.54445 | 1 | 0.075749 | -0.90699 | 0.91015 | -0.74374
|  |  | -6.5 | 0.32733 | 0.0077257 | 0.28058 | 0.5575 | 0.60791 | 0.046752 | -0.54977 | 1 | 0.076906 | -0.90436 | 0.90762 | -0.74298
|  |  | -6 | 0.32562 | 0.007782 | 0.2793 | 0.55207 | 0.60492 | 0.046325 | -0.54428 | 1 | 0.07658 | -0.89976 | 0.90301 | -0.74294
|  |  | -5.5 | 0.33759 | 0.0078822 | 0.29083 | 0.57111 | 0.62842 | 0.046753 | -0.56323 | 1 | 0.074398 | -0.89626 | 0.89934 | -0.74399
|  |  | -5 | 0.33795 | 0.0080348 | 0.29004 | 0.57227 | 0.62799 | 0.047913 | -0.56423 | 1 | 0.076295 | -0.89847 | 0.9017 | -0.74304
|  |  | -4.5 | 0.3407 | 0.0080205 | 0.29383 | 0.58002 | 0.63453 | 0.046875 | -0.572 | 1 | 0.073874 | -0.90146 | 0.90448 | -0.74451
|  |  | -4 | 0.3443 | 0.0081139 | 0.29608 | 0.58197 | 0.64038 | 0.048218 | -0.57386 | 1 | 0.075295 | -0.89612 | 0.89928 | -0.74348
|  |  | -3.5 | 0.34778 | 0.0079814 | 0.30316 | 0.59088 | 0.65094 | 0.044616 | -0.5829 | 1 | 0.06854 | -0.89547 | 0.89809 | -0.7472
|  |  | -3 | 0.34491 | 0.0080987 | 0.30023 | 0.59314 | 0.64515 | 0.044678 | -0.58504 | 1 | 0.069253 | -0.90684 | 0.90948 | -0.74729
|  |  | -2.5 | 0.34454 | 0.0084801 | 0.29602 | 0.58197 | 0.64057 | 0.048523 | -0.57349 | 1 | 0.07575 | -0.89529 | 0.89849 | -0.74319
|  |  | -2 | 0.35034 | 0.0081692 | 0.30493 | 0.59644 | 0.65528 | 0.04541 | -0.58827 | 1 | 0.069299 | -0.89774 | 0.90041 | -0.74688
|  |  | -1.5 | 0.33875 | 0.0077887 | 0.29553 | 0.57746 | 0.63428 | 0.043213 | -0.56967 | 1 | 0.068129 | -0.89813 | 0.90071 | -0.74754
|  |  | -1 | 0.34326 | 0.0083094 | 0.29895 | 0.5821 | 0.64222 | 0.044312 | -0.57379 | 1 | 0.068999 | -0.89345 | 0.89611 | -0.74686
|  |  | -0.5 | 0.36621 | 0.0079518 | 0.31714 | 0.61487 | 0.68335 | 0.049072 | -0.60692 | 1 | 0.071811 | -0.88815 | 0.89105 | -0.74506
|  |  | 0 | 0.34345 | 0.0083352 | 0.29901 | 0.58649 | 0.64246 | 0.044434 | -0.57816 | 1 | 0.069162 | -0.89991 | 0.90256 | -0.74705
|  |  | 0.5 | 0.35651 | 0.0080824 | 0.31159 | 0.61139 | 0.66809 | 0.044922 | -0.60331 | 1 | 0.067239 | -0.90303 | 0.90553 | -0.74824
|  |  | 1 | 0.3244 | 0.0082789 | 0.27979 | 0.55024 | 0.60419 | 0.044617 | -0.54196 | 1 | 0.073846 | -0.897 | 0.90003 | -0.74433
|  |  | 1.5 | 0.36536 | 0.0085078 | 0.31098 | 0.61023 | 0.67633 | 0.054382 | -0.60173 | 1 | 0.080407 | -0.88969 | 0.89331 | -0.74033
|  |  | 2 | 0.33002 | 0.0084773 | 0.2804 | 0.56177 | 0.61042 | 0.049622 | -0.55329 | 1 | 0.081292 | -0.90642 | 0.91006 | -0.74068
|  |  | 2.5 | 0.3432 | 0.0080519 | 0.29578 | 0.57917 | 0.63898 | 0.047425 | -0.57111 | 1 | 0.07422 | -0.89379 | 0.89687 | -0.74397
|  |  | 3 | 0.33563 | 0.0082675 | 0.28546 | 0.56507 | 0.6211 | 0.050171 | -0.5568 | 1 | 0.080778 | -0.89648 | 0.90011 | -0.74047
|  |  | 3.5 | 0.32935 | 0.0077772 | 0.28449 | 0.56592 | 0.61383 | 0.044861 | -0.55814 | 1 | 0.073083 | -0.90927 | 0.91221 | -0.7453
|  |  | 4 | 0.32575 | 0.0077677 | 0.28143 | 0.55256 | 0.60718 | 0.044312 | -0.54479 | 1 | 0.07298 | -0.89724 | 0.9002 | -0.74482
|  |  | 4.5 | 0.3216 | 0.0077267 | 0.27631 | 0.54688 | 0.5979 | 0.045288 | -0.53915 | 1 | 0.075745 | -0.90174 | 0.90491 | -0.7435
|  |  | 5 | 0.31928 | 0.0075379 | 0.27643 | 0.54328 | 0.59571 | 0.042847 | -0.53574 | 1 | 0.071926 | -0.89933 | 0.90221 | -0.74549
|  |  | 5.5 | 0.31329 | 0.0072727 | 0.27429 | 0.53858 | 0.58759 | 0.039001 | -0.53131 | 1 | 0.066375 | -0.90421 | 0.90665 | -0.74876
|  |  | 6 | 0.31146 | 0.0073223 | 0.26752 | 0.53229 | 0.57898 | 0.043945 | -0.52497 | 1 | 0.0759 | -0.90671 | 0.90988 | -0.74364
|  |  | 6.5 | 0.30353 | 0.0071106 | 0.26428 | 0.52076 | 0.56781 | 0.039246 | -0.51365 | 1 | 0.069118 | -0.9046 | 0.90724 | -0.74727
|  |  | 7 | 0.29175 | 0.0070515 | 0.25055 | 0.49738 | 0.5423 | 0.041199 | -0.49033 | 1 | 0.075971 | -0.90416 | 0.90734 | -0.74348
|  |  | 7.5 | 0.297 | 0.0069981 | 0.25397 | 0.50684 | 0.55097 | 0.04303 | -0.49984 | 1 | 0.078099 | -0.90721 | 0.91056 | -0.74246
|  |  | 8 | 0.28571 | 0.0066872 | 0.24664 | 0.48822 | 0.53235 | 0.039063 | -0.48153 | 1 | 0.073378 | -0.90454 | 0.90752 | -0.74493
|  |  | 8.5 | 0.26477 | 0.0065995 | 0.22931 | 0.45734 | 0.49408 | 0.035462 | -0.45074 | 1 | 0.071774 | -0.91227 | 0.91509 | -0.74614
|  |  | 9 | 0.27313 | 0.006424 | 0.23358 | 0.46399 | 0.50672 | 0.039552 | -0.45757 | 1 | 0.078056 | -0.903 | 0.90637 | -0.74229
|  |  | 9.5 | 0.26874 | 0.0062428 | 0.23059 | 0.46558 | 0.49933 | 0.038148 | -0.45934 | 1 | 0.076398 | -0.9199 | 0.92307 | -0.74397
|  |  | 10 | 0.24994 | 0.0059414 | 0.21594 | 0.43543 | 0.46588 | 0.033997 | -0.42949 | 1 | 0.072973 | -0.92187 | 0.92476 | -0.7459
|  |  | 10.5 | 0.24976 | 0.0060196 | 0.21509 | 0.43451 | 0.46485 | 0.034668 | -0.42849 | 1 | 0.07458 | -0.92179 | 0.9248 | -0.74503
|  |  | 11 | 0.23956 | 0.0057164 | 0.20685 | 0.41815 | 0.44641 | 0.032715 | -0.41244 | 1 | 0.073284 | -0.92389 | 0.92679 | -0.74582
|  |  | 11.5 | 0.22522 | 0.0054627 | 0.19782 | 0.39997 | 0.42304 | 0.027405 | -0.3945 | 1 | 0.064782 | -0.93255 | 0.9348 | -0.75072
|  |  | 12 | 0.22943 | 0.0051489 | 0.20099 | 0.40357 | 0.43042 | 0.028442 | -0.39842 | 1 | 0.066079 | -0.92564 | 0.928 | -0.74976
|  |  | 12.5 | 0.2359 | 0.0050822 | 0.20508 | 0.41113 | 0.44098 | 0.030823 | -0.40605 | 1 | 0.069896 | -0.92079 | 0.92344 | -0.74752
|  |  | 13 | 0.21063 | 0.0047999 | 0.18341 | 0.36902 | 0.39404 | 0.027221 | -0.36422 | 1 | 0.069081 | -0.92431 | 0.92689 | -0.7481
|  |  | 13.5 | 0.2005 | 0.0045729 | 0.17664 | 0.35694 | 0.37714 | 0.023864 | -0.35236 | 1 | 0.063277 | -0.93431 | 0.93645 | -0.75159
|  |  | 14 | 0.19482 | 0.0044842 | 0.16919 | 0.34021 | 0.36401 | 0.025635 | -0.33573 | 1 | 0.070423 | -0.92229 | 0.92498 | -0.74729
|  |  | 14.5 | 0.18616 | 0.004241 | 0.16327 | 0.32465 | 0.34943 | 0.022888 | -0.32041 | 1 | 0.065501 | -0.91695 | 0.91928 | -0.74974
|  |  | 15 | 0.17713 | 0.0040942 | 0.15527 | 0.31311 | 0.3324 | 0.021851 | -0.30902 | 1 | 0.065737 | -0.92966 | 0.93198 | -0.7501
|  |  | 15.5 | 0.17334 | 0.0039587 | 0.15235 | 0.30634 | 0.32569 | 0.020996 | -0.30238 | 1 | 0.064467 | -0.92843 | 0.93067 | -0.75074
|  |  | 16 | 0.17365 | 0.0040436 | 0.15192 | 0.30823 | 0.32556 | 0.021729 | -0.30419 | 1 | 0.066743 | -0.93434 | 0.93672 | -0.74974
|  |  | 16.5 | 0.17102 | 0.0037623 | 0.15112 | 0.30353 | 0.32215 | 0.019897 | -0.29977 | 1 | 0.061764 | -0.93054 | 0.93258 | -0.75226
|  |  | 17 | 0.16467 | 0.0035219 | 0.14313 | 0.28998 | 0.3078 | 0.021546 | -0.28646 | 1 | 0.07 | -0.93066 | 0.93329 | -0.74786
|  |  | 17.5 | 0.14301 | 0.0034332 | 0.12524 | 0.25488 | 0.26825 | 0.017761 | -0.25145 | 1 | 0.06621 | -0.93737 | 0.93971 | -0.75014
|  |  | 18 | 0.14551 | 0.0032816 | 0.12714 | 0.25946 | 0.27265 | 0.018372 | -0.25618 | 1 | 0.067384 | -0.93961 | 0.94202 | -0.7496
|  |  | 18.5 | 0.13867 | 0.0032187 | 0.12073 | 0.24646 | 0.2594 | 0.017945 | -0.24324 | 1 | 0.069179 | -0.93771 | 0.94026 | -0.74858
|  |  | 19 | 0.13599 | 0.0031366 | 0.11731 | 0.23963 | 0.2533 | 0.018677 | -0.23649 | 1 | 0.073736 | -0.93364 | 0.93655 | -0.74599
|  |  | 19.5 | 0.12512 | 0.0029173 | 0.10877 | 0.22394 | 0.23389 | 0.016358 | -0.22102 | 1 | 0.069939 | -0.94499 | 0.94757 | -0.74846
|  |  | 20 | 0.12927 | 0.0028544 | 0.11322 | 0.23206 | 0.24249 | 0.016052 | -0.2292 | 1 | 0.066195 | -0.94519 | 0.9475 | -0.75044
|  |  | 20.5 | 0.11322 | 0.0026722 | 0.099369 | 0.20404 | 0.21259 | 0.013852 | -0.20137 | 1 | 0.065156 | -0.94721 | 0.94945 | -0.75106
|  |  | 21 | 0.11621 | 0.0025663 | 0.10248 | 0.20746 | 0.21869 | 0.013732 | -0.20489 | 1 | 0.062792 | -0.93691 | 0.93901 | -0.75194
|  |  | 21.5 | 0.1084 | 0.0024767 | 0.095551 | 0.19653 | 0.20395 | 0.012848 | -0.19406 | 1 | 0.062996 | -0.95149 | 0.95358 | -0.75234
|  |  | 22 | 0.10376 | 0.0024109 | 0.090332 | 0.18482 | 0.19409 | 0.013428 | -0.1824 | 1 | 0.069181 | -0.93978 | 0.94232 | -0.74866
|  |  | 22.5 | 0.10071 | 0.0022965 | 0.08667 | 0.17853 | 0.18738 | 0.014038 | -0.17623 | 1 | 0.074916 | -0.94052 | 0.9435 | -0.74566
|  |  | 23 | 0.097821 | 0.0022564 | 0.084626 | 0.17578 | 0.18245 | 0.013195 | -0.17353 | 1 | 0.072323 | -0.9511 | 0.95385 | -0.74745
|  |  | 23.5 | 0.094262 | 0.0020781 | 0.082306 | 0.16937 | 0.17657 | 0.011955 | -0.1673 | 1 | 0.067709 | -0.94749 | 0.9499 | -0.74973
|  |  | 24 | 0.088372 | 0.0020304 | 0.076661 | 0.15985 | 0.16503 | 0.011711 | -0.15782 | 1 | 0.070963 | -0.95631 | 0.95894 | -0.74836
|  |  | 24.5 | 0.084065 | 0.0019035 | 0.07367 | 0.15308 | 0.15773 | 0.010395 | -0.15117 | 1 | 0.065902 | -0.9584 | 0.96067 | -0.75107
|  |  | 25 | 0.080491 | 0.0017967 | 0.070832 | 0.14679 | 0.15132 | 0.0096588 | -0.14499 | 1 | 0.063829 | -0.95818 | 0.9603 | -0.75214
|  |  | 25.5 | 0.075081 | 0.0017297 | 0.065857 | 0.13727 | 0.14094 | 0.0092239 | -0.13554 | 1 | 0.065446 | -0.96169 | 0.96391 | -0.75142
|  |  | 26 | 0.073395 | 0.00166 | 0.064606 | 0.13379 | 0.138 | 0.0087891 | -0.13213 | 1 | 0.063688 | -0.95745 | 0.95957 | -0.75219
|  |  | 26.5 | 0.071583 | 0.0016115 | 0.063202 | 0.13123 | 0.13479 | 0.0083809 | -0.12962 | 1 | 0.062179 | -0.96164 | 0.96365 | -0.75311
|  |  | 27 | 0.067219 | 0.0015149 | 0.059052 | 0.12073 | 0.12627 | 0.0081675 | -0.11921 | 1 | 0.064682 | -0.9441 | 0.94632 | -0.7512
|  |  | 27.5 | 0.068512 | 0.0014939 | 0.059998 | 0.12262 | 0.12851 | 0.0085146 | -0.12113 | 1 | 0.066256 | -0.94254 | 0.94487 | -0.75031
|  |  | 28 | 0.064122 | 0.0014403 | 0.056061 | 0.1181 | 0.12018 | 0.0080607 | -0.11666 | 1 | 0.06707 | -0.97072 | 0.97303 | -0.75091
|  |  | 28.5 | 0.058026 | 0.0013591 | 0.050812 | 0.10651 | 0.10884 | 0.0072136 | -0.10515 | 1 | 0.066279 | -0.9661 | 0.96837 | -0.75115
|  |  | 29 | 0.058449 | 0.0013477 | 0.050537 | 0.10596 | 0.10899 | 0.0079117 | -0.10461 | 1 | 0.072594 | -0.95985 | 0.96259 | -0.74765
|  |  | 29.5 | 0.053868 | 0.0012823 | 0.046517 | 0.097485 | 0.10038 | 0.0073509 | -0.096203 | 1 | 0.073228 | -0.95834 | 0.96114 | -0.74727
|  |  | 30 | 0.048931 | 0.0012043 | 0.043255 | 0.090424 | 0.092187 | 0.0056763 | -0.08922 | 1 | 0.061574 | -0.96782 | 0.96978 | -0.75363
|  |  | 30.5 | 0.05077 | 0.0011876 | 0.044497 | 0.09372 | 0.095267 | 0.0062732 | -0.092532 | 1 | 0.065849 | -0.9713 | 0.97353 | -0.75155
|  |  | 31 | 0.049527 | 0.0011333 | 0.043162 | 0.090172 | 0.092688 | 0.0063649 | -0.089039 | 1 | 0.06867 | -0.96063 | 0.96308 | -0.74972
|  |  | 31.5 | 0.048165 | 0.0010784 | 0.042431 | 0.089039 | 0.090596 | 0.0057335 | -0.087961 | 1 | 0.063287 | -0.97092 | 0.97298 | -0.75285
|  |  | 32 | 0.045376 | 0.0010395 | 0.039759 | 0.083317 | 0.085135 | 0.0056172 | -0.082278 | 1 | 0.06598 | -0.96644 | 0.96869 | -0.75132
|  |  | 32.5 | 0.043778 | 0.0010062 | 0.038508 | 0.080693 | 0.082285 | 0.00527 | -0.079687 | 1 | 0.064045 | -0.96842 | 0.97053 | -0.75238
|  |  | 33 | 0.03888 | 0.00094891 | 0.034396 | 0.071846 | 0.073275 | 0.0044841 | -0.070898 | 1 | 0.061195 | -0.96755 | 0.96949 | -0.75382
|  |  | 33.5 | 0.039135 | 0.00094843 | 0.034285 | 0.072186 | 0.07342 | 0.0048504 | -0.071238 | 1 | 0.066064 | -0.97027 | 0.97252 | -0.75141
|  |  | 34 | 0.036778 | 0.00089145 | 0.032757 | 0.068791 | 0.069535 | 0.0040207 | -0.067899 | 1 | 0.057823 | -0.97648 | 0.97819 | -0.75582
|  |  | 34.5 | 0.035252 | 0.00084258 | 0.030922 | 0.064469 | 0.066174 | 0.0043298 | -0.063626 | 1 | 0.065431 | -0.9615 | 0.96372 | -0.75143
|  |  | 35 | 0.03648 | 0.00081408 | 0.032507 | 0.068692 | 0.068987 | 0.003973 | -0.067878 | 1 | 0.05759 | -0.98391 | 0.9856 | -0.75617
|  |  | 35.5 | 0.034249 | 0.00081528 | 0.029915 | 0.062718 | 0.064164 | 0.0043336 | -0.061903 | 1 | 0.06754 | -0.96476 | 0.96712 | -0.75045
|  |  | 36 | 0.028938 | 0.00077987 | 0.025536 | 0.052895 | 0.054474 | 0.0034027 | -0.052115 | 1 | 0.062465 | -0.95669 | 0.95873 | -0.7528
|  |  | 36.5 | 0.028767 | 0.00072992 | 0.025349 | 0.053078 | 0.054116 | 0.003418 | -0.052348 | 1 | 0.063161 | -0.96734 | 0.9694 | -0.7528
|  |  | 37 | 0.027699 | 0.00072658 | 0.024283 | 0.051659 | 0.051981 | 0.003416 | -0.050932 | 1 | 0.065716 | -0.97982 | 0.98202 | -0.75191
|  |  | 37.5 | 0.024963 | 0.0007093 | 0.022352 | 0.047127 | 0.047316 | 0.0026112 | -0.046418 | 1 | 0.055187 | -0.98102 | 0.98257 | -0.7573
|  |  | 38 | 0.025597 | 0.00066662 | 0.022711 | 0.047161 | 0.048308 | 0.0028858 | -0.046495 | 1 | 0.059738 | -0.96247 | 0.96432 | -0.7544
|  |  | 38.5 | 0.028359 | 0.00069964 | 0.024561 | 0.051735 | 0.05292 | 0.0037976 | -0.051036 | 1 | 0.071762 | -0.9644 | 0.96706 | -0.74826
|  |  | 39 | 0.027535 | 0.00066018 | 0.024107 | 0.051102 | 0.051642 | 0.0034275 | -0.050442 | 1 | 0.066371 | -0.97676 | 0.97902 | -0.75148
|  |  | 39.5 | 0.022461 | 0.00061322 | 0.019669 | 0.041245 | 0.04213 | 0.0027924 | -0.040631 | 1 | 0.066281 | -0.96444 | 0.96671 | -0.75109
|  |  | 40 | 0.027779 | 0.00071061 | 0.023901 | 0.051079 | 0.05168 | 0.0038776 | -0.050368 | 1 | 0.075031 | -0.97462 | 0.97751 | -0.74698
|  |  | 40.5 | 0.01923 | 0.00062287 | 0.016911 | 0.035523 | 0.036141 | 0.0023193 | -0.0349 | 1 | 0.064174 | -0.96567 | 0.9678 | -0.75222
|  |  | 41 | 0.0261 | 0.00053871 | 0.022251 | 0.048001 | 0.048351 | 0.0038491 | -0.047462 | 1 | 0.079607 | -0.9816 | 0.98482 | -0.74494
|  |  | 41.5 | 0.026051 | 0.00039506 | 0.02236 | 0.047737 | 0.048411 | 0.0036907 | -0.047342 | 1 | 0.076237 | -0.97793 | 0.9809 | -0.7465
|  |  | 42 | 0.01759 | 0.00022829 | 0.014851 | 0.031254 | 0.03244 | 0.002739 | -0.031026 | 1 | 0.084432 | -0.95639 | 0.96011 | -0.74137
|  |  | 42.5 | 0.0084353 | 0.0001453 | 0.0072441 | 0.015457 | 0.015679 | 0.0011912 | -0.015312 | 1 | 0.07597 | -0.97656 | 0.97951 | -0.74658
|  |  | 43 | 0.0042072 | 7.9967e-05 | 0.0035186 | 0.007339 | 0.0077258 | 0.00068856 | -0.0072591 | 1 | 0.089125 | -0.93959 | 0.94381 | -0.73811
|  |  | 43.5 | 0.0019341 | 5.132e-05 | 0.0015826 | 0.0033541 | 0.0035167 | 0.00035143 | -0.0033028 | 1 | 0.099932 | -0.93917 | 0.94447 | -0.7324
|  |  | 44 | 0.00094784 | 3.7559e-05 | 0.00080228 | 0.0017204 | 0.0017501 | 0.00014556 | -0.0016829 | 1 | 0.083169 | -0.96158 | 0.96517 | -0.74226
|  |  | 44.5 | 0.00059009 | 2.8432e-05 | 0.0005591 | 0.0011177 | 0.0011492 | 3.0995e-05 | -0.0010893 | 1 | 0.026971 | -0.94787 | 0.94825 | -0.77117
|  |  | 45 | 0.0004046 | 2.9899e-05 | 0.00038302 | 0.00074345 | 0.00078762 | 2.1577e-05 | -0.00071355 | 1 | 0.027395 | -0.90596 | 0.90638 | -0.77028
|  |  | 45.5 | 0.0003624 | 2.601e-05 | 0.00033438 | 0.00063867 | 0.00069678 | 2.8014e-05 | -0.00061266 | 1 | 0.040205 | -0.87927 | 0.88019 | -0.76255
|  |  | 46 | 0.00038499 | 2.4699e-05 | 0.00032213 | 0.00064981 | 0.00070712 | 6.2853e-05 | -0.00062511 | 1 | 0.088886 | -0.88403 | 0.88848 | -0.73529
|  |  | 46.5 | 0.0002206 | 2.6606e-05 | 0.00019923 | 0.00040424 | 0.00041983 | 2.1369e-05 | -0.00037764 | 1 | 0.0509 | -0.8995 | 0.90094 | -0.75714
|  |  | 47 | 0.00012577 | 2.1331e-05 | 0.00012648 | 0.00024504 | 0.00025225 | -7.15e-07 | -0.00022371 | 1 | -0.0028345 | -0.88685 | 0.88685 | 0.7838
|  |  | 47.5 | 0.00015634 | 2.3663e-05 | 0.00012937 | 0.00027293 | 0.00028572 | 2.6971e-05 | -0.00024927 | 1 | 0.094398 | -0.87243 | 0.87753 | -0.73151
|  |  | 48 | 0.00016076 | 2.5124e-05 | 0.00018501 | 0.00034744 | 0.00034577 | -2.4259e-05 | -0.00032231 | 1 | -0.07016 | -0.93216 | 0.9348 | 0.74784
|  |  | 48.5 | 0.00018639 | 2.6219e-05 | 0.00012985 | 0.00028861 | 0.00031623 | 5.6535e-05 | -0.00026239 | 1 | 0.17878 | -0.82973 | 0.84877 | -0.67929
|  |  | 49 | 0.00012589 | 2.0743e-05 | 0.00010106 | 0.00020582 | 0.00022695 | 2.4826e-05 | -0.00018507 | 1 | 0.10939 | -0.81549 | 0.8228 | -0.71873
|  |  | 49.5 | 0.00011557 | 1.9392e-05 | 0.00011781 | 0.00021368 | 0.00023338 | -2.235e-06 | -0.00019429 | 1 | -0.0095765 | -0.8325 | 0.83256 | 0.77965
|  |  | 50 | 8.5958e-05 | 2.3495e-05 | 8.1271e-05 | 0.0001508 | 0.00016723 | 4.687e-06 | -0.00012731 | 1 | 0.028027 | -0.76127 | 0.76178 | -0.767
|  |  | 50.5 | 0.00012463 | 2.4717e-05 | 9.6769e-05 | 0.00019086 | 0.0002214 | 2.7865e-05 | -0.00016614 | 1 | 0.12586 | -0.75039 | 0.76087 | -0.70231
|  |  | 51 | 0.00010538 | 3.2742e-05 | 7.7218e-05 | 0.00015301 | 0.0001826 | 2.8163e-05 | -0.00012026 | 1 | 0.15423 | -0.65862 | 0.67644 | -0.67038
|  |  | 51.5 | 0.00012928 | 2.3052e-05 | 8.7798e-05 | 0.00018317 | 0.00021708 | 4.1485e-05 | -0.00016011 | 1 | 0.1911 | -0.73758 | 0.76193 | -0.65864
|  |  | 52 | 0.00011551 | 2.2207e-05 | 9.2596e-05 | 0.0001902 | 0.00020811 | 2.2918e-05 | -0.00016799 | 1 | 0.11012 | -0.80723 | 0.8147 | -0.71761
|  |  | 52.5 | 7.5758e-05 | 1.8727e-05 | 6.8546e-05 | 0.00013232 | 0.0001443 | 7.212e-06 | -0.0001136 | 1 | 0.049978 | -0.7872 | 0.78878 | -0.7537
|  |  | 53 | 7.9394e-05 | 2.3741e-05 | 7.2867e-05 | 0.00012261 | 0.00015226 | 6.527e-06 | -9.8867e-05 | 1 | 0.042867 | -0.64933 | 0.65074 | -0.75244
|  |  | 53.5 | 0.00010306 | 3.0842e-05 | 6.935e-05 | 0.00013906 | 0.00017241 | 3.3707e-05 | -0.00010822 | 1 | 0.19551 | -0.62768 | 0.65743 | -0.63442
|  |  | 54 | 7.847e-05 | 2.4572e-05 | 6.6221e-05 | 0.0001201 | 0.00014469 | 1.2249e-05 | -9.5532e-05 | 1 | 0.084656 | -0.66025 | 0.66565 | -0.72164
|  |  | 54.5 | 6.7294e-05 | 2.0895e-05 | 7.4566e-05 | 0.00012219 | 0.00014186 | -7.272e-06 | -0.00010129 | 1 | -0.051262 | -0.71405 | 0.71589 | 0.74956
|  |  | 55 | 0.00010166 | 2.2609e-05 | 0.00011256 | 0.00019181 | 0.00021422 | -1.0908e-05 | -0.0001692 | 1 | -0.05092 | -0.78984 | 0.79148 | 0.75321
|  |  | 55.5 | 7.6056e-05 | 2.2907e-05 | 5.7668e-05 | 0.00012487 | 0.00013372 | 1.8388e-05 | -0.00010197 | 1 | 0.13751 | -0.76251 | 0.77481 | -0.69619
|  |  | 56 | 9.5845e-05 | 2.2125e-05 | 9.8378e-05 | 0.00017697 | 0.00019422 | -2.533e-06 | -0.00015484 | 1 | -0.013042 | -0.79724 | 0.79734 | 0.77722
|  |  | 56.5 | 8.8603e-05 | 1.7531e-05 | 6.6579e-05 | 0.0001334 | 0.00015518 | 2.2024e-05 | -0.00011587 | 1 | 0.14192 | -0.74664 | 0.76001 | -0.69148
|  |  | 57 | 9.5189e-05 | 2.246e-05 | 8.2821e-05 | 0.0001502 | 0.00017801 | 1.2368e-05 | -0.00012774 | 1 | 0.069479 | -0.71763 | 0.72098 | -0.73714
|  |  | 57.5 | 7.3553e-05 | 1.9241e-05 | 5.6505e-05 | 0.00011152 | 0.00013006 | 1.7048e-05 | -9.228e-05 | 1 | 0.13108 | -0.70953 | 0.72154 | -0.69406
|  |  | 58 | 5.132e-05 | 1.8642e-05 | 5.0605e-05 | 9.3278e-05 | 0.00010192 | 7.15e-07 | -7.4636e-05 | 1 | 0.007015 | -0.73226 | 0.7323 | -0.78061
|  |  | 58.5 | 9.5428e-05 | 1.8746e-05 | 0.00010285 | 0.00017795 | 0.00019828 | -7.42e-06 | -0.0001592 | 1 | -0.037423 | -0.80295 | 0.80382 | 0.76211
|  |  | 59 | 3.8125e-05 | 1.5065e-05 | 3.5158e-05 | 6.1453e-05 | 7.3283e-05 | 2.967e-06 | -4.6388e-05 | 1 | 0.040487 | -0.633 | 0.63429 | -0.75346
|  |  | 59.5 | 4.6762e-05 | 1.9126e-05 | 6.5938e-05 | 8.7977e-05 | 0.0001127 | -1.9176e-05 | -6.8851e-05 | 1 | -0.17015 | -0.61092 | 0.63418 | 0.64958
|  |  | 60 | 9.8832e-05 | 2.3935e-05 | 6.7361e-05 | 0.00014275 | 0.00016619 | 3.1471e-05 | -0.00011882 | 1 | 0.18936 | -0.71495 | 0.7396 | -0.65594
Note on A13: FO: The setup to measure the optical power output vs the angle is composed by a motorized rotation stage (central hub) with an arm. The detection sits at the end of the arm and is composed by a Glan-Thompson polarizer mounted in a motorized rotation stage. An aperture with 6.5 mm is monted after the polarized to reject the undesired polarization. It is followed by a slit with a 1 mm width and a detector. The distance between central hub and the slit is ~573mm (1 mm slit = 0,1º). The detector is connected to a lock-in amplifer SR830. The laser is installed in the diode holder with the laser diode window centred with the central hub. A wave generator is used to drive the lase diode controller with a sinosoidal wave of amplitude 'A' and offset of 'A/2'. The power meter was used to adjust the average power. The laser was measured with the diode die oriented roughly with the horizontal and vertical of the laboratory to measure the angular dispertion in both orientations. Polariation was measured at 0º (horizontal), 45º, 90º and 135º.