## Information
Laser Diode Information
Datasheet Information
Brand: | Roithner Lasers
Part Number: | S6705MG
Package Size: | 5.6 mm (TO56), Flat window
Photodiode: | Integrated
Wavelength: | 670 | nm
Optical Power: | 7 | mW
Threshold Current: | 24 | mA
Operating Current: | 34 | mA
Slope Efficiency: | 0.5 | W/A
Beam Divergence (FWHM)
parallel: | 9 | deg
perpendicular: | 32 | deg

## OpticalPower_vs_Current_20ºC
Optical Power Output vs Current
Measurment Information: |  |  | Measured Data: | [7 empty] | Calculated Data:
|  |  | * Power value was not calibrated | [7 empty] | * Thorlabs LDC205C current accuracy = +/-0.5 mA
PowerMeter: | Thorlabs PM320E |  | ** flutuation of the measured value shown in the device display. | [7 empty] | ** Thorlabs LDC205C photocurrent accuracy = +/- 10 uA
PowerMeter Detector: | Thorlabs S142C | (Set to 673 nm) | [8 empty] | *** Thorlabs S142C = +/- 3%; Thorlabs PM320E = +/- 0.2%
Laser Diode Controller: | Thorlabs LDC205C
Temperature Controller: | Throlabs TED200C |  | Current |  | Optical Power* |  | Photodiode Current |  | Optical Power |  | Current |  | Optical Power |  | Photodiode Current |  | Optical Power
Laser Diode Holder: | Thorlabs TCLDM9 |  | I_LD | (mA) | P_LD | (mW) | I_PD | (mA) | mW |  | I_LD | Error* | P_LD | Error | I_PD | Error** | mW | Error***
Measurment Notes |  |  | min** | max** | min | max | min | max | min | max | (ºC) | (+/- ºC) | (mW) | (+/- mW) | (mA) | (+/- mA) | (mW) | (+/- mW)
|  |  | 23.98 | 23.98 | 0.256 | 0.256 | 0.025 | 0.025 | 0.494 | 0.495 | 23.98 | 0.5 | 0.256 |  | 0.025 | 0.01 | 0.4945 | 0.015368
Fixed LD house Temp.: | 19.995 | ºC | 25.02 | 25.02 | 0.569 | 0.57 | 0.055 | 0.055 | 1.104 | 1.104 | 25.02 | 0.5 | 0.5695 |  | 0.055 | 0.01 | 1.104 | 0.033194
|  |  | 26.11 | 26.11 | 0.897 | 0.898 | 0.086 | 0.087 | 1.735 | 1.736 | 26.11 | 0.5 | 0.8975 |  | 0.0865 | 0.0105 | 1.7355 | 0.052681
|  |  | 27.1 | 27.11 | 1.204 | 1.204 | 0.116 | 0.116 | 2.323 | 2.325 | 27.105 | 0.505 | 1.204 |  | 0.116 | 0.01 | 2.324 | 0.070875
|  |  | 27.99 | 28 | 1.482 | 1.483 | 0.143 | 0.143 | 2.859 | 2.86 | 27.995 | 0.505 | 1.4825 |  | 0.143 | 0.01 | 2.8595 | 0.086475
|  |  | 29 | 29 | 1.796 | 1.797 | 0.173 | 0.173 | 3.467 | 3.469 | 29 | 0.5 | 1.7965 |  | 0.173 | 0.01 | 3.468 | 0.10527
|  |  | 30.02 | 30.03 | 2.122 | 2.123 | 0.204 | 0.204 | 4.098 | 4.1 | 30.025 | 0.505 | 2.1225 |  | 0.204 | 0.01 | 4.099 | 0.12424
|  |  | 31.01 | 31.02 | 2.44 | 2.441 | 0.235 | 0.235 | 4.709 | 4.71 | 31.015 | 0.505 | 2.4405 |  | 0.235 | 0.01 | 4.7095 | 0.1421
|  |  | 32.1 | 32.1 | 2.787 | 2.788 | 0.268 | 0.268 | 5.376 | 5.378 | 32.1 | 0.5 | 2.7875 |  | 0.268 | 0.01 | 5.377 | 0.16267
|  |  | 33 | 33 | 3.079 | 3.08 | 0.296 | 0.296 | 5.938 | 5.939 | 33 | 0.5 | 3.0795 |  | 0.296 | 0.01 | 5.9385 | 0.17905
|  |  | 34.03 | 34.03 | 3.414 | 3.415 | 0.329 | 0.329 | 6.58 | 6.583 | 34.03 | 0.5 | 3.4145 |  | 0.329 | 0.01 | 6.5815 | 0.19938
|  |  | 35.01 | 35.01 | 3.735 | 3.736 | 0.36 | 0.36 | 7.2 | 7.202 | 35.01 | 0.5 | 3.7355 |  | 0.36 | 0.01 | 7.201 | 0.21751
|  |  | 36.02 | 36.02 | 4.062 | 4.063 | 0.391 | 0.391 | 7.834 | 7.835 | 36.02 | 0.5 | 4.0625 |  | 0.391 | 0.01 | 7.8345 | 0.23606
|  |  | 37.04 | 37.04 | 4.393 | 4.394 | 0.423 | 0.423 | 8.48 | 8.483 | 37.04 | 0.5 | 4.3935 |  | 0.423 | 0.01 | 8.4815 | 0.25651
Note on A10: FO: The setup to measure the optical power output of the laser is composed of the devices listed previously. The diode and the detector had to be as close as possible to measure the full laser beam. The front plate of the Laser Diode Holder was removed and the same did happen to the SM1-Threaded part at the entrance of the Detector. Removing these parts helped reduzing the distance between the diode and the detector. The Detector was pressed again the laser diode holder to measure the full beam without using optics.

## OpticalPower_vs_Current_25º
Optical Power Output vs Current
Measurment Information: |  |  | Measured Data: | [7 empty] | Calculated Data:
|  |  | * Power value was not calibrated | [7 empty] | * Thorlabs LDC205C current accuracy = +/-0.5 mA
PowerMeter: | Thorlabs PM320E |  | ** flutuation of the measured value shown in the device display. | [7 empty] | ** Thorlabs LDC205C photocurrent accuracy = +/- 10 uA
PowerMeter Detector: | Thorlabs S142C | (Set to 673 nm) | [8 empty] | *** Thorlabs S142C = +/- 3%; Thorlabs PM320E = +/- 0.2%
Laser Diode Controller: | Thorlabs LDC205C
Temperature Controller: | Throlabs TED200C |  | Current |  | Optical Power* |  | Photodiode Current |  | Optical Power |  | Current |  | Optical Power |  | Photodiode Current |  | Optical Power
Laser Diode Holder: | Thorlabs TCLDM9 |  | I_LD | (mA) | P_LD | (mW) | I_PD | (mA) | mW |  | I_LD | Error* | P_LD | Error | I_PD | Error** | mW | Error***
Measurment Notes |  |  | min** | max** | min | max | min | max | min | max | (ºC) | (+/- ºC) | (mW) | (+/- mW) | (mA) | (+/- mA) | (mW) | (+/- mW)
|  |  | 24 | 24 | 0.047 | 0.048 | 0.005 | 0.005 | 0.1023 | 0.1025 | 24 | 0.5 | 0.0475 |  | 0.005 | 0.01 | 0.1024 | 0.0031788
Fixed LD house Temp.: | 25 | ºC | 25 | 25 | 0.117 | 0.117 | 0.011 | 0.011 | 0.246 | 0.248 | 25 | 0.5 | 0.117 |  | 0.011 | 0.01 | 0.247 | 0.0084264
|  |  | 26.01 | 26.01 | 0.392 | 0.392 | 0.038 | 0.038 | 0.819 | 0.821 | 26.01 | 0.5 | 0.392 |  | 0.038 | 0.01 | 0.82 | 0.025655
|  |  | 27.01 | 27.01 | 0.689 | 0.689 | 0.066 | 0.066 | 1.441 | 1.443 | 27.01 | 0.5 | 0.689 |  | 0.066 | 0.01 | 1.442 | 0.044356
|  |  | 28.07 | 28.08 | 1.015 | 1.015 | 0.097 | 0.098 | 2.114 | 2.116 | 28.075 | 0.505 | 1.015 |  | 0.0975 | 0.0105 | 2.115 | 0.064591
|  |  | 29.04 | 29.04 | 1.309 | 1.309 | 0.126 | 0.126 | 2.726 | 2.728 | 29.04 | 0.5 | 1.309 |  | 0.126 | 0.01 | 2.727 | 0.082992
|  |  | 30.02 | 30.03 | 1.614 | 1.615 | 0.155 | 0.155 | 3.364 | 3.365 | 30.025 | 0.505 | 1.6145 |  | 0.155 | 0.01 | 3.3645 | 0.10166
|  |  | 31.05 | 31.05 | 1.935 | 1.935 | 0.186 | 0.186 | 4.027 | 4.029 | 31.05 | 0.5 | 1.935 |  | 0.186 | 0.01 | 4.028 | 0.12211
|  |  | 32.01 | 32.01 | 2.232 | 2.233 | 0.215 | 0.215 | 4.646 | 4.648 | 32.01 | 0.5 | 2.2325 |  | 0.215 | 0.01 | 4.647 | 0.14072
|  |  | 33.01 | 33.02 | 2.551 | 2.552 | 0.246 | 0.246 | 5.309 | 5.315 | 33.015 | 0.505 | 2.5515 |  | 0.246 | 0.01 | 5.312 | 0.16271
|  |  | 34.02 | 34.03 | 2.867 | 2.867 | 0.276 | 0.276 | 5.974 | 5.976 | 34.025 | 0.505 | 2.867 |  | 0.276 | 0.01 | 5.975 | 0.18065
|  |  | 34.99 | 34.99 | 3.175 | 3.175 | 0.306 | 0.306 | 6.616 | 6.618 | 34.99 | 0.5 | 3.175 |  | 0.306 | 0.01 | 6.617 | 0.19995
Note on A10: FO: The setup to measure the optical power output of the laser is composed of the devices listed previously. The diode and the detector had to be as close as possible to measure the full laser beam. The front plate of the Laser Diode Holder was removed and the same did happen to the SM1-Threaded part at the entrance of the Detector. Removing these parts helped reduzing the distance between the diode and the detector. The Detector was pressed again the laser diode holder to measure the full beam without using optics.

## OpticalPower_vs_Temp
Optical Power Output vs Current
Measurment Information: |  |  | Measured Data: | [7 empty] | Calculated Data:
|  |  | * Power value was not calibrated | [7 empty] | * Thorlabs LDC205C current accuracy = +/-0.5 mA
PowerMeter: | Thorlabs PM320E |  | ** flutuation of the measured value shown in the device display. | [7 empty] | ** Thorlabs LDC205C photocurrent accuracy = +/- 10 uA
PowerMeter Detector: | Thorlabs S142C | (Set to 673 nm) | [8 empty] | *** Thorlabs S142C = +/- 3%; Thorlabs PM320E = +/- 0.2%
Laser Diode Controller: | Thorlabs LDC205C
Temperature Controller: | Throlabs TED200C |  | Temperature |  | Optical Power* |  | Photodiode Current |  | Optical Power |  | Current |  | Optical Power |  | Photodiode Current |  | Optical Power
Laser Diode Holder: | Thorlabs TCLDM9 |  | T | (ºC) | P_LD | (mW) | I_PD | (mA) | mW |  | I_LD | Error* | P_LD | Error | I_PD | Error** | mW | Error***
Measurment Notes |  |  | min** | max** | min | max | min | max | min | max | (ºC) | (+/- ºC) | (mW) | (+/- mW) | (mA) | (+/- mA) | (mW) | (+/- mW)
|  |  | 25 | 25 | 2.114 | 2.115 | 0.204 | 0.204 | 4.095 | 4.096 | 25 | 0.5 | 2.1145 |  | 0.204 | 0.01 | 4.0955 | 0.12364
Fixed LD current: | 31.01 | mA | 23 | 23 | 2.24 | 2.241 | 0.216 | 0.216 | 4.335 | 4.336 | 23 | 0.5 | 2.2405 |  | 0.216 | 0.01 | 4.3355 | 0.13085
|  |  | 20 | 20 | 2.428 | 2.429 | 0.234 | 0.234 | 4.683 | 4.686 | 20 | 0.5 | 2.4285 |  | 0.234 | 0.01 | 4.6845 | 0.14235
|  |  | 16.99 | 16.99 | 2.612 | 2.613 | 0.252 | 0.252 | 5.025 | 5.028 | 16.99 | 0.5 | 2.6125 |  | 0.252 | 0.01 | 5.0265 | 0.15263
|  |  | 15.01 | 15.01 | 2.738 | 2.739 | 0.264 | 0.264 | 5.25 | 5.251 | 15.01 | 0.5 | 2.7385 |  | 0.264 | 0.01 | 5.2505 | 0.15836
|  |  | 12 | 12 | 2.916 | 2.917 | 0.281 | 0.281 | 5.585 | 5.586 | 12 | 0.5 | 2.9165 |  | 0.281 | 0.01 | 5.5855 | 0.16844
|  |  | 10 | 10 | 3.036 | 3.037 | 0.292 | 0.292 | 5.799 | 5.799 | 10 | 0.5 | 3.0365 |  | 0.292 | 0.01 | 5.799 | 0.17436
|  |  | 6.99 | 7 | 3.218 | 3.219 | 0.31 | 0.31 | 6.123 | 6.126 | 6.995 | 0.505 | 3.2185 |  | 0.31 | 0.01 | 6.1245 | 0.18564
|  |  | 25 | 25 | 2.147 | 2.148 | 0.207 | 0.207 | 4.152 | 4.153 | 25 | 0.5 | 2.1475 |  | 0.207 | 0.01 | 4.1525 | 0.12535
Note on A10: FO: The setup to measure the optical power output of the laser is composed of the devices listed previously. The diode and the detector had to be as close as possible to measure the full laser beam. The front plate of the Laser Diode Holder was removed and the same did happen to the SM1-Threaded part at the entrance of the Detector. Removing these parts helped reduzing the distance between the diode and the detector. The Detector was pressed again the laser diode holder to measure the full beam without using optics.